## Slide 1: HB Field Mapping
Steven Lassiter
Eric Sun
Mike Fowler

## Slide 2: Index
Mapping setup
Bore central line map
Stray field map @ 1200A
Stray field map @ 4000A
Coil movement
Off midplane stray field Maps @3500A
Integral stray field along exit beam line & Tosca yoke modifications.

## Slide 3: HB Stray Field Mapping Setup
Central field tube
Stray field mapping table and probe holder

## Slide 4: HB Magnet - Yoke & Coil
Point #27 is used as midpoint

## Slide 5: Notes
Lakeshore Hall probes used to measure the By field (vertical component).
Stray field Hall probe alignment of ~1mm. (surveyed positions, hand aligned probe).
Central field Hall probe alignment of the order 3mm in Z. (Tape measure).
Spacing between measured points is 1”.
Fields not corrected for current variations. Current was very stable throughout.
Field mapping table is not perfectly aligned to midplane axis of magnet.

## Slide 6: HB Central Field vs Tosca
Point #27 | Measured | Tosca @ point | Difference | Error %
Current | By [kG] | By [kG] | Gauss | 
1200 | 8.079 | 8.010 | 69.3 | 0.86
2000 | 13.416 | 13.342 | 73.9 | 0.55
3000 | 19.907 | 19.879 | 27.6 | 0.14
3500 | 22.932 | 22.933 | -0.9 | 0.00
4000 | 25.695 | 25.771 | -75.6 | -0.29
Chart: HB B vs I Curve
| 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1200 | 1500 | 2000 | 2500 | 3000 | 3100 | 3200 | 3300 | 3400 | 3500 | 3600 | 3700 | 3800 | 3900 | 4000
Measured | 0.6745 | 1.3468 | 2.0196 | 2.6927 | 3.3662 | 4.0398 | 4.7145 | 5.3874 | 6.0598 | 6.7314 | 8.079 | 10.0788 | 13.4156 | 16.6925 | 19.9068 | 20.5188 | 21.1367 | 21.7462 | 22.3439 | 22.9319 | 23.4987 | 24.0597 | 24.6115 | 25.1563 | 25.6952
Tosca @ probe 27 | 8.0097 | 13.3417 | 19.8792 | 22.9329 | 25.7708
Error % | 0.85728 | 0.5512 | 0.1385 | -0.0041299 | -0.29433
Difference | 69.2597 | 73.9476 | 27.5703 | -0.94705 | -75.629

## Slide 7
Chart: HB Central Line
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58
1200A | 0.24439 | 0.31358 | 0.38999 | 0.51378 | 0.66102 | 0.8639 | 1.1248 | 1.5112 | 1.945 | 0 | 3.1412 | 3.8496 | 4.5526 | 5.3175 | 6.0199 | 6.5904 | 7.0969 | 7.4333 | 7.6814 | 7.8411 | 7.9426 | 8.0037 | 8.0436 | 8.066 | 8.0741 | 8.0787 | 8.079 | 8.0755 | 8.0696 | 8.0575 | 8.0405 | 8.0145 | 7.9726 | 7.9057 | 7.8079 | 7.653 | 7.4171 | 7.0902 | 6.6365 | 6.1045 | 5.4578 | 4.7656 | 4.0728 | 3.4242 | 2.8009 | 2.2534 | 1.7876 | 1.4025 | 1.0917 | 0.85312 | 0.66785 | 0.51975 | 0.41063 | 0.32291 | 0.2591 | 0.20843 | 0.16546 | 0
2000A | 0.40188 | 0.50989 | 0.62266 | 0.85067 | 1.0782 | 1.4114 | 1.87 | 2.4324 | 3.1853 | 0 | 5.1918 | 6.3323 | 7.5064 | 8.6552 | 9.8853 | 10.8621 | 11.6759 | 12.2907 | 12.7048 | 12.9727 | 13.1548 | 13.2743 | 13.3455 | 13.3896 | 13.4077 | 13.4148 | 13.4156 | 13.4096 | 13.397 | 13.3774 | 13.343 | 13.2964 | 13.2183 | 13.0968 | 12.9107 | 12.6281 | 12.2326 | 11.6606 | 10.9123 | 9.9875 | 8.961 | 7.8074 | 6.6674 | 5.628 | 4.6151 | 3.6988 | 2.9523 | 2.2832 | 1.797 | 1.4022 | 1.0837 | 0.84873 | 0.66114 | 0.52884 | 0.42076 | 0.33806 | 0.27121 | 0
3000A | 0.57333 | 0.74524 | 0.92988 | 1.2024 | 1.5567 | 2.0692 | 2.6557 | 3.6032 | 4.6771 | 0 | 7.4975 | 9.0955 | 10.9435 | 12.5845 | 14.2932 | 15.7241 | 16.9219 | 17.8044 | 18.4958 | 18.9853 | 19.3518 | 19.5807 | 19.7386 | 19.8366 | 19.8738 | 19.9006 | 19.9068 | 19.8949 | 19.8652 | 19.8131 | 19.7311 | 19.6079 | 19.4448 | 19.1752 | 18.835 | 18.3368 | 17.6561 | 16.7551 | 15.703 | 14.3211 | 12.8242 | 11.2714 | 9.6202 | 8.0915 | 6.6593 | 5.3284 | 4.2056 | 3.2879 | 2.5889 | 2.0141 | 1.5698 | 1.2256 | 0.96348 | 0.75687 | 0.60371 | 0.47917 | 0.3871 | 0
4000A | 0.70681 | 0.91614 | 1.1668 | 1.5311 | 1.9609 | 2.5893 | 3.4305 | 4.5987 | 5.9616 | 0 | 9.6115 | 11.6382 | 13.9068 | 16.0539 | 18.1996 | 20.0492 | 21.4978 | 22.7454 | 23.6385 | 24.2738 | 24.7868 | 25.1372 | 25.3858 | 25.5637 | 25.6368 | 25.6851 | 25.695 | 25.6662 | 25.5999 | 25.5014 | 25.3389 | 25.1329 | 24.8436 | 24.4472 | 23.9458 | 23.2843 | 22.3516 | 21.2356 | 19.7999 | 18.183 | 16.289 | 14.2809 | 12.3172 | 10.2843 | 8.4102 | 6.8352 | 5.3814 | 4.1773 | 3.2764 | 2.5464 | 1.9612 | 1.5234 | 1.1933 | 0.93689 | 0.74509 | 0.58912 | 0.46784
Tosca | 0.63604 | 0.81033 | 1.0409 | 1.3462 | 1.7523 | 2.2927 | 3.0049 | 3.9412 | 5.1318 | 6.6099 | 8.3552 | 10.3125 | 12.3723 | 14.4092 | 16.3062 | 17.9676 | 19.3436 | 20.4241 | 21.2291 | 21.8168 | 22.2309 | 22.5122 | 22.7015 | 22.8227 | 22.8944 | 22.9289 | 22.9329 | 22.9131 | 22.8628 | 22.7827 | 22.6635 | 22.4935 | 22.2565 | 21.9274 | 21.4774 | 20.869 | 20.0559 | 19.0098 | 17.708 | 16.1672 | 14.4374 | 12.5958 | 10.7322 | 8.9364 | 7.2906 | 5.8434 | 4.616 | 3.6143 | 2.8158 | 2.1899 | 1.7052 | 1.3285 | 1.0449 | 0.82635 | 0.65704 | 0.52688 | 0.42502 | 0.34581 | 0.2834
Tosca | 0.69877 | 0.89315 | 1.1507 | 1.4922 | 1.9473 | 2.5536 | 3.3534 | 4.4056 | 5.7441 | 7.4056 | 9.3662 | 11.5623 | 13.8686 | 16.1442 | 18.2586 | 20.1081 | 21.6416 | 22.8505 | 23.7592 | 24.431 | 24.9127 | 25.2472 | 25.4779 | 25.6296 | 25.7215 | 25.7663 | 25.7708 | 25.7422 | 25.6749 | 25.5682 | 25.4128 | 25.1976 | 24.9055 | 24.511 | 23.9831 | 23.2838 | 22.3625 | 21.1899 | 19.7401 | 18.0298 | 16.1102 | 14.0642 | 11.989 | 9.9853 | 8.1453 | 6.5252 | 5.1507 | 4.0277 | 3.133 | 2.4322 | 1.8898 | 1.4687 | 1.1524 | 0.90875 | 0.72038 | 0.57584 | 0.46295 | 0.37534 | 0.30647
Tosca | 0.23842 | 0.3022 | 0.38632 | 0.49737 | 0.64475 | 0.84031 | 1.0979 | 1.4361 | 1.8662 | 2.401 | 3.0344 | 3.7479 | 4.5029 | 5.2522 | 5.9484 | 6.5486 | 7.0278 | 7.3813 | 7.6235 | 7.7827 | 7.8826 | 7.9419 | 7.9772 | 7.9969 | 8.0069 | 8.0106 | 8.0097 | 8.0069 | 7.9986 | 7.9868 | 7.9692 | 7.942 | 7.9009 | 7.8362 | 7.7366 | 7.5831 | 7.3524 | 7.0207 | 6.5747 | 6.0171 | 5.3732 | 4.6794 | 3.9779 | 3.3057 | 2.6936 | 2.1584 | 1.7058 | 1.3383 | 1.0452 | 0.81542 | 0.63731 | 0.49887 | 0.39397 | 0.3131 | 0.25023 | 0.20173 | 0.16363 | 0.13388 | 0.11035
Tosca | 0.39271 | 0.498 | 0.63692 | 0.82035 | 1.0638 | 1.387 | 1.8126 | 2.3715 | 3.0821 | 3.9651 | 5.0102 | 6.1864 | 7.4302 | 8.6646 | 9.8139 | 10.8111 | 11.6165 | 12.2206 | 12.6424 | 12.9248 | 13.1049 | 13.2134 | 13.2788 | 13.316 | 13.3351 | 13.3426 | 13.3417 | 13.3366 | 13.3216 | 13.2997 | 13.2665 | 13.2152 | 13.1375 | 13.0163 | 12.8318 | 12.5525 | 12.1417 | 11.5664 | 10.8098 | 9.8807 | 8.8193 | 7.6823 | 6.534 | 5.433 | 4.4292 | 3.5503 | 2.8063 | 2.2016 | 1.7192 | 1.3409 | 1.0477 | 0.81972 | 0.6471 | 0.514 | 0.41059 | 0.33081 | 0.26819 | 0.21931 | 0.18065
Tosca | 0.56372 | 0.71654 | 0.91843 | 1.1854 | 1.5402 | 2.0117 | 2.6328 | 3.4489 | 4.4863 | 5.7746 | 7.2969 | 9.0063 | 10.8087 | 12.5947 | 14.2601 | 15.7178 | 16.9196 | 17.8537 | 18.5386 | 19.0266 | 19.36 | 19.5778 | 19.7188 | 19.8055 | 19.8547 | 19.8773 | 19.8792 | 19.8668 | 19.8324 | 19.7784 | 19.6969 | 19.5759 | 19.4011 | 19.1488 | 18.7907 | 18.2896 | 17.602 | 16.6996 | 15.5624 | 14.2066 | 12.6806 | 11.0559 | 9.4147 | 7.8363 | 6.3926 | 5.125 | 4.0505 | 3.1746 | 2.4761 | 1.9284 | 1.504 | 1.1741 | 0.92513 | 0.73315 | 0.58423 | 0.46957 | 0.37971 | 0.30971 | 0.25446
3500A | 0.64069 | 0.8229 | 1.0478 | 1.3746 | 1.7685 | 2.3449 | 3.0648 | 4.1013 | 5.3237 | 8.6301 | 10.6221 | 12.514 | 14.4354 | 16.3608 | 17.9498 | 19.3135 | 20.4094 | 21.1761 | 21.7571 | 22.1703 | 22.4772 | 22.6818 | 22.823 | 22.8847 | 22.9229 | 22.9319 | 22.9116 | 22.868 | 22.7844 | 22.6658 | 22.5028 | 22.2602 | 21.9246 | 21.463 | 20.8736 | 20.0833 | 19.023 | 17.8042 | 16.3029 | 14.5772 | 12.8211 | 11.0579 | 9.2332 | 7.5827 | 6.1126 | 4.8528 | 3.7521 | 2.9287 | 2.2869 | 1.7645 | 1.3842 | 1.0682 | 0.85188 | 0.6803 | 0.53651 | 0.43152
Coil F | 27 | 0 | 0 | 27 | 27
Yoke | 27 | 0 | 0 | 27 | 27

## Slide 8
Chart: HB 1200A Pos
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58
1200A | 0.24439 | 0.31358 | 0.38999 | 0.51378 | 0.66102 | 0.8639 | 1.1248 | 1.5112 | 1.945 | 0 | 3.1412 | 3.8496 | 4.5526 | 5.3175 | 6.0199 | 6.5904 | 7.0969 | 7.4333 | 7.6814 | 7.8411 | 7.9426 | 8.0037 | 8.0436 | 8.066 | 8.0741 | 8.0787 | 8.079 | 8.0755 | 8.0696 | 8.0575 | 8.0405 | 8.0145 | 7.9726 | 7.9057 | 7.8079 | 7.653 | 7.4171 | 7.0902 | 6.6365 | 6.1045 | 5.4578 | 4.7656 | 4.0728 | 3.4242 | 2.8009 | 2.2534 | 1.7876 | 1.4025 | 1.0917 | 0.85312 | 0.66785 | 0.51975 | 0.41063 | 0.32291 | 0.2591 | 0.20843 | 0.16546 | 0
Tosca | 0.23842 | 0.3022 | 0.38632 | 0.49737 | 0.64475 | 0.84031 | 1.0979 | 1.4361 | 1.8662 | 2.401 | 3.0344 | 3.7479 | 4.5029 | 5.2522 | 5.9484 | 6.5486 | 7.0278 | 7.3813 | 7.6235 | 7.7827 | 7.8826 | 7.9419 | 7.9772 | 7.9969 | 8.0069 | 8.0106 | 8.0097 | 8.0069 | 7.9986 | 7.9868 | 7.9692 | 7.942 | 7.9009 | 7.8362 | 7.7366 | 7.5831 | 7.3524 | 7.0207 | 6.5747 | 6.0171 | 5.3732 | 4.6794 | 3.9779 | 3.3057 | 2.6936 | 2.1584 | 1.7058 | 1.3383 | 1.0452 | 0.81542 | 0.63731 | 0.49887 | 0.39397 | 0.3131 | 0.25023 | 0.20173 | 0.16363 | 0.13388 | 0.11035
Difference | 5.9658 | 11.3798 | 3.6682 | 16.4057 | 16.2711 | 23.5871 | 26.8937 | 75.1234 | 78.7931 | 106.8505 | 101.7182 | 49.626 | 65.2701 | 71.5101 | 41.8066 | 69.1476 | 51.9064 | 57.8657 | 58.347 | 59.9451 | 61.8082 | 66.413 | 69.0296 | 67.2035 | 68.1292 | 69.2597 | 68.6643 | 70.9748 | 70.6556 | 71.2575 | 72.5161 | 71.637 | 69.4863 | 71.3082 | 69.8883 | 64.7605 | 69.4895 | 61.7768 | 87.3346 | 84.6276 | 86.1976 | 94.8784 | 118.5466 | 107.3064 | 95.0446 | 81.8017 | 64.2983 | 46.5923 | 37.6979 | 30.5421 | 20.8847 | 16.6578 | 9.8144 | 8.8671 | 6.7049 | 1.8269

## Slide 9
Chart: HB 2000A Pos
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58
2000A | 0.40188 | 0.50989 | 0.62266 | 0.85067 | 1.0782 | 1.4114 | 1.87 | 2.4324 | 3.1853 | 0 | 5.1918 | 6.3323 | 7.5064 | 8.6552 | 9.8853 | 10.8621 | 11.6759 | 12.2907 | 12.7048 | 12.9727 | 13.1548 | 13.2743 | 13.3455 | 13.3896 | 13.4077 | 13.4148 | 13.4156 | 13.4096 | 13.397 | 13.3774 | 13.343 | 13.2964 | 13.2183 | 13.0968 | 12.9107 | 12.6281 | 12.2326 | 11.6606 | 10.9123 | 9.9875 | 8.961 | 7.8074 | 6.6674 | 5.628 | 4.6151 | 3.6988 | 2.9523 | 2.2832 | 1.797 | 1.4022 | 1.0837 | 0.84873 | 0.66114 | 0.52884 | 0.42076 | 0.33806 | 0.27121 | 0
Tosca | 0.39271 | 0.498 | 0.63692 | 0.82035 | 1.0638 | 1.387 | 1.8126 | 2.3715 | 3.0821 | 3.9651 | 5.0102 | 6.1864 | 7.4302 | 8.6646 | 9.8139 | 10.8111 | 11.6165 | 12.2206 | 12.6424 | 12.9248 | 13.1049 | 13.2134 | 13.2788 | 13.316 | 13.3351 | 13.3426 | 13.3417 | 13.3366 | 13.3216 | 13.2997 | 13.2665 | 13.2152 | 13.1375 | 13.0163 | 12.8318 | 12.5525 | 12.1417 | 11.5664 | 10.8098 | 9.8807 | 8.8193 | 7.6823 | 6.534 | 5.433 | 4.4292 | 3.5503 | 2.8063 | 2.2016 | 1.7192 | 1.3409 | 1.0477 | 0.81972 | 0.6471 | 0.514 | 0.41059 | 0.33081 | 0.26819 | 0.21931 | 0.18065
Difference | 9.1722 | 11.8917 | -14.255 | 30.3176 | 14.3207 | 24.3864 | 57.3406 | 60.9509 | 103.2109 | 181.669 | 145.8999 | 76.2658 | -9.4002 | 71.4108 | 51.0088 | 59.3743 | 70.0209 | 62.431 | 47.9601 | 49.8845 | 60.9022 | 66.6659 | 73.6589 | 72.5637 | 72.1918 | 73.9476 | 72.9946 | 75.3893 | 77.7142 | 76.4436 | 81.2174 | 80.7242 | 80.52 | 78.9236 | 75.6403 | 90.8429 | 94.1863 | 102.502 | 106.7423 | 141.7378 | 125.1693 | 133.3332 | 195.0901 | 185.9767 | 148.5075 | 145.9959 | 81.5637 | 77.8671 | 61.3013 | 35.9923 | 29.0103 | 14.0374 | 14.8357 | 10.1738 | 7.2468 | 3.0174

## Slide 10
Chart: HB 3000A Pos
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58
3000A | 0.57333 | 0.74524 | 0.92988 | 1.2024 | 1.5567 | 2.0692 | 2.6557 | 3.6032 | 4.6771 | 0 | 7.4975 | 9.0955 | 10.9435 | 12.5845 | 14.2932 | 15.7241 | 16.9219 | 17.8044 | 18.4958 | 18.9853 | 19.3518 | 19.5807 | 19.7386 | 19.8366 | 19.8738 | 19.9006 | 19.9068 | 19.8949 | 19.8652 | 19.8131 | 19.7311 | 19.6079 | 19.4448 | 19.1752 | 18.835 | 18.3368 | 17.6561 | 16.7551 | 15.703 | 14.3211 | 12.8242 | 11.2714 | 9.6202 | 8.0915 | 6.6593 | 5.3284 | 4.2056 | 3.2879 | 2.5889 | 2.0141 | 1.5698 | 1.2256 | 0.96348 | 0.75687 | 0.60371 | 0.47917 | 0.3871 | 0
Tosca | 0.56372 | 0.71654 | 0.91843 | 1.1854 | 1.5402 | 2.0117 | 2.6328 | 3.4489 | 4.4863 | 5.7746 | 7.2969 | 9.0063 | 10.8087 | 12.5947 | 14.2601 | 15.7178 | 16.9196 | 17.8537 | 18.5386 | 19.0266 | 19.36 | 19.5778 | 19.7188 | 19.8055 | 19.8547 | 19.8773 | 19.8792 | 19.8668 | 19.8324 | 19.7784 | 19.6969 | 19.5759 | 19.4011 | 19.1488 | 18.7907 | 18.2896 | 17.602 | 16.6996 | 15.5624 | 14.2066 | 12.6806 | 11.0559 | 9.4147 | 7.8363 | 6.3926 | 5.125 | 4.0505 | 3.1746 | 2.4761 | 1.9284 | 1.504 | 1.1741 | 0.92513 | 0.73315 | 0.58423 | 0.46957 | 0.37971 | 0.30971 | 0.25446
Difference | 9.6093 | 28.7022 | 11.4504 | 16.986 | 16.4811 | 57.5149 | 22.8287 | 154.2718 | 190.7657 | 200.5983 | 89.1717 | 134.7872 | -10.2059 | 33.0987 | 6.2771 | 2.2197 | -49.3202 | -42.8346 | -41.2456 | -8.2017 | 2.85 | 19.8174 | 31.1197 | 19.1884 | 23.336 | 27.5703 | 28.1313 | 32.8043 | 34.6751 | 34.147 | 32.0717 | 43.6549 | 26.3777 | 44.302 | 47.1833 | 54.0862 | 55.4866 | 140.6459 | 114.4696 | 143.6268 | 215.5212 | 205.5432 | 255.2249 | 266.7398 | 203.4142 | 155.1565 | 113.254 | 112.7896 | 85.6377 | 65.8123 | 51.4613 | 38.3526 | 23.7223 | 19.4778 | 9.6032 | 7.3889

## Slide 11
Chart: HB 3500A Pos
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58
3500A | 0.64069 | 0.8229 | 1.0478 | 1.3746 | 1.7685 | 2.3449 | 3.0648 | 4.1013 | 5.3237 | 8.6301 | 10.6221 | 12.514 | 14.4354 | 16.3608 | 17.9498 | 19.3135 | 20.4094 | 21.1761 | 21.7571 | 22.1703 | 22.4772 | 22.6818 | 22.823 | 22.8847 | 22.9229 | 22.9319 | 22.9116 | 22.868 | 22.7844 | 22.6658 | 22.5028 | 22.2602 | 21.9246 | 21.463 | 20.8736 | 20.0833 | 19.023 | 17.8042 | 16.3029 | 14.5772 | 12.8211 | 11.0579 | 9.2332 | 7.5827 | 6.1126 | 4.8528 | 3.7521 | 2.9287 | 2.2869 | 1.7645 | 1.3842 | 1.0682 | 0.85188 | 0.6803 | 0.53651 | 0.43152
Tosca | 0.63604 | 0.81033 | 1.0409 | 1.3462 | 1.7523 | 2.2927 | 3.0049 | 3.9412 | 5.1318 | 6.6099 | 8.3552 | 10.3125 | 12.3723 | 14.4092 | 16.3062 | 17.9676 | 19.3436 | 20.4241 | 21.2291 | 21.8168 | 22.2309 | 22.5122 | 22.7015 | 22.8227 | 22.8944 | 22.9289 | 22.9329 | 22.9131 | 22.8628 | 22.7827 | 22.6635 | 22.4935 | 22.2565 | 21.9274 | 21.4774 | 20.869 | 20.0559 | 19.0098 | 17.708 | 16.1672 | 14.4374 | 12.5958 | 10.7322 | 8.9364 | 7.2906 | 5.8434 | 4.616 | 3.6143 | 2.8158 | 2.1899 | 1.7052 | 1.3285 | 1.0449 | 0.82635 | 0.65704 | 0.52688 | 0.42502 | 0.34581 | 0.2834
Difference | 4.6534 | 12.5658 | 6.9362 | 28.4451 | 16.1597 | 52.2063 | 59.9233 | 160.0755 | 191.8868 | 274.9391 | 309.5732 | 141.7818 | 26.1625 | 54.5657 | -17.8484 | -30.1119 | -14.7165 | -52.9827 | -59.6786 | -60.5621 | -35.0108 | -19.6398 | 0.29297 | -9.7678 | -6.0192 | -0.94705 | -1.5231 | 5.1258 | 1.724 | 2.3129 | 9.2108 | 3.708 | -2.711 | -14.3281 | 4.5658 | 27.3634 | 13.2451 | 96.182 | 135.6789 | 139.8374 | 225.2689 | 325.7489 | 296.7423 | 292.0496 | 269.2331 | 236.8138 | 137.7774 | 112.9798 | 97.0247 | 59.3273 | 55.6624 | 23.3085 | 25.5317 | 23.2574 | 9.635 | 6.4966

## Slide 12
Chart: HB 4000A Pos
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60
Tosca | 0.69877 | 0.89315 | 1.1507 | 1.4922 | 1.9473 | 2.5536 | 3.3534 | 4.4056 | 5.7441 | 7.4056 | 9.3662 | 11.5623 | 13.8686 | 16.1442 | 18.2586 | 20.1081 | 21.6416 | 22.8505 | 23.7592 | 24.431 | 24.9127 | 25.2472 | 25.4779 | 25.6296 | 25.7215 | 25.7663 | 25.7708 | 25.7422 | 25.6749 | 25.5682 | 25.4128 | 25.1976 | 24.9055 | 24.511 | 23.9831 | 23.2838 | 22.3625 | 21.1899 | 19.7401 | 18.0298 | 16.1102 | 14.0642 | 11.989 | 9.9853 | 8.1453 | 6.5252 | 5.1507 | 4.0277 | 3.133 | 2.4322 | 1.8898 | 1.4687 | 1.1524 | 0.90875 | 0.72038 | 0.57584 | 0.46295 | 0.37534 | 0.30647
4000A | 0.70681 | 0.91614 | 1.1668 | 1.5311 | 1.9609 | 2.5893 | 3.4305 | 4.5987 | 5.9616 | 0 | 9.6115 | 11.6382 | 13.9068 | 16.0539 | 18.1996 | 20.0492 | 21.4978 | 22.7454 | 23.6385 | 24.2738 | 24.7868 | 25.1372 | 25.3858 | 25.5637 | 25.6368 | 25.6851 | 25.695 | 25.6662 | 25.5999 | 25.5014 | 25.3389 | 25.1329 | 24.8436 | 24.4472 | 23.9458 | 23.2843 | 22.3516 | 21.2356 | 19.7999 | 18.183 | 16.289 | 14.2809 | 12.3172 | 10.2843 | 8.4102 | 6.8352 | 5.3814 | 4.1773 | 3.2764 | 2.5464 | 1.9612 | 1.5234 | 1.1933 | 0.93689 | 0.74509 | 0.58912 | 0.46784
Difference | 8.0499 | 16.0646 | 38.8174 | 13.6093 | 35.6616 | 77.1125 | 193.0666 | 217.501 | 245.2551 | 75.9302 | 38.1955 | -90.2774 | -58.9938 | -58.9616 | -143.7872 | -105.0142 | -120.6689 | -157.2452 | -125.8717 | -110.0224 | -92.039 | -65.8305 | -84.6625 | -81.2378 | -75.809 | -74.9878 | -66.826 | -73.9379 | -64.7682 | -61.8894 | -63.7539 | -37.3527 | 0.56943 | -10.8996 | 45.6762 | 59.8 | 178.7908 | 216.7413 | 328.1893 | 299.043 | 264.8911 | 310.0334 | 230.7254 | 143.4005 | 71.3325 | 40.9514 | 24.7054 | 4.8904

## Slide 13: Stray Field map vs Beam line @ 5.5° Row 1 nearly coincides with 5.5° beam line

## Slide 14: Tosca Field Map @ 1,200A

## Slide 15: Stray By Fields at 1,200A
Chart: HB Stray Field Measurment at 1200A
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Row 6 | -0.11179 | -0.11359 | -0.11031 | -0.10041 | -0.08277 | -0.05726 | -0.02535 | 0.01022 | 0.04595 | 0.07835 | 0.10467 | 0.12374 | 0.13564 | 0.14155 | 0.14318 | 0.14178 | 0.13983 | 0.13585 | 0.13259 | 0.12975 | 0.12754 | 0.12525 | 0.12283 | 0.1207 | 0.11925 | 0.11835 | 0.11803 | 0.11806 | 0.11845 | 0.1191 | 0.11966 | 0.11985 | 0.11886 | 0.11591 | 0.11026 | 0.10113 | 0.08835 | 0.07207 | 0.05291 | 0.03291 | 0.01297 | -0.00533 | -0.02081 | -0.03282 | -0.0413 | -0.04631 | -0.0491
Row 5 | -0.14375 | -0.14944 | -0.14846 | -0.1381 | -0.11581 | -0.0812 | -0.0356 | 0.01656 | 0.06909 | 0.11652 | 0.1538 | 0.17894 | 0.19271 | 0.19771 | 0.19659 | 0.19184 | 0.18653 | 0.18011 | 0.17486 | 0.17169 | 0.16955 | 0.16665 | 0.16202 | 0.15832 | 0.15602 | 0.15477 | 0.15435 | 0.1546 | 0.15547 | 0.15688 | 0.15857 | 0.16003 | 0.16038 | 0.1583 | 0.15196 | 0.13997 | 0.12258 | 0.09939 | 0.07253 | 0.04371 | 0.0154 | -0.00998 | -0.03073 | -0.04615 | -0.0562 | -0.06181 | -0.06369
Row 4 | -0.18674 | -0.19978 | -0.20432 | -0.19564 | -0.16982 | -0.12373 | -0.05799 | 0.02019 | 0.10093 | 0.17266 | 0.22712 | 0.26096 | 0.27646 | 0.27821 | 0.27183 | 0.2616 | 0.25077 | 0.24075 | 0.23267 | 0.22725 | 0.2256 | 0.22224 | 0.21437 | 0.20856 | 0.20532 | 0.20362 | 0.20312 | 0.2037 | 0.20501 | 0.20762 | 0.21096 | 0.2146 | 0.21773 | 0.2176 | 0.2117 | 0.19765 | 0.1749 | 0.14087 | 0.10114 | 0.05892 | 0.01783 | -0.01779 | -0.04569 | -0.06537 | -0.07707 | -0.0823 | -0.08276
Row 3 | -0.24083 | -0.26701 | -0.28311 | -0.28175 | -0.25601 | -0.19891 | -0.10605 | 0.01219 | 0.13786 | 0.2493 | 0.33044 | 0.37589 | 0.39162 | 0.38609 | 0.37065 | 0.35157 | 0.33361 | 0.31827 | 0.3063 | 0.29843 | 0.29507 | 0.29005 | 0.28091 | 0.27449 | 0.27043 | 0.26823 | 0.26824 | 0.26879 | 0.27087 | 0.27483 | 0.28142 | 0.28915 | 0.29605 | 0.30002 | 0.29546 | 0.2784 | 0.2459 | 0.19857 | 0.13829 | 0.07573 | 0.01641 | -0.03523 | -0.07193 | -0.09592 | -0.10809 | -0.11139 | -0.10859
Row 2 | -0.30746 | -0.35516 | -0.39288 | -0.41147 | -0.39874 | -0.33954 | -0.21855 | -0.04239 | 0.15791 | 0.33906 | 0.46567 | 0.52854 | 0.5414 | 0.52265 | 0.49172 | 0.45919 | 0.43132 | 0.41028 | 0.39361 | 0.35553 | 0.35123 | 0.3491 | 0.34957 | 0.35386 | 0.35673 | 0.36068 | 0.36961 | 0.3841 | 0.39977 | 0.41259 | 0.41411 | 0.39473 | 0.34918 | 0.27703 | 0.18568 | 0.08906 | 0.00117 | -0.06901 | -0.11597 | -0.14247 | -0.15193 | -0.1503 | -0.14168
Row 1 | -0.38615 | -0.46093 | -0.52998 | -0.58731 | -0.62213 | -0.6007 | -0.48396 | -0.25137 | 0.06699 | 0.37489 | 0.58759 | 0.68145 | 0.68533 | 0.64352 | 0.5897 | 0.54132 | 0.50394 | 0.47838 | 0.4618 | 0.42977 | 0.42683 | 0.4252 | 0.42662 | 0.4342 | 0.44683 | 0.45665 | 0.46983 | 0.49291 | 0.52526 | 0.55791 | 0.57355 | 0.55491 | 0.48982 | 0.37825 | 0.23289 | 0.08532 | -0.04374 | -0.13688 | -0.19075 | -0.21383 | -0.21424 | -0.20254 | -0.18381

## Slide 16: Row 1 @ 1,200A
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -398.2851 | -470.905 | -544.7506 | -613.9313 | -658.4438 | -652.7708 | -562.8342 | -358.5496 | -62.0559 | 247.2577 | 492.1569 | 627.7608 | 661.1316 | 635.9268 | 593.8346 | 554.4172 | 523.327 | 499.4859 | 482.3252 | 447.072 | 445.8535 | 446.6381 | 449.4433 | 455.2496 | 465.0723 | 479.2515 | 497.9364 | 519.4515 | 538.7005 | 545.3863 | 523.7918 | 460.5589 | 354.5226 | 215.9137 | 65.317 | -66.2108 | -162.8061 | -221.5294 | -245.7222 | -246.2954 | -232.6723 | -211.4705 | -188.0068
Measured | -386.15 | -460.93 | -529.98 | -587.31 | -622.13 | -600.7 | -483.96 | -251.37 | 66.99 | 374.89 | 587.59 | 681.45 | 685.33 | 643.52 | 589.7 | 541.32 | 503.94 | 478.38 | 461.8 | 429.77 | 426.83 | 425.2 | 426.62 | 434.2 | 446.83 | 456.65 | 469.83 | 492.91 | 525.26 | 557.91 | 573.55 | 554.91 | 489.82 | 378.25 | 232.89 | 85.32 | -43.74 | -136.88 | -190.75 | -213.83 | -214.24 | -202.54 | -183.81
Difference | 12.1351 | 9.975 | 14.7706 | 26.6213 | 36.3138 | 52.0708 | 78.8742 | 107.1796 | 129.0459 | 127.6323 | 95.4331 | 53.6892 | 24.1984 | 7.5932 | -4.1346 | -13.0972 | -19.387 | -21.1059 | -20.5252 | -17.302 | -19.0235 | -21.4381 | -22.8233 | -21.0496 | -18.2423 | -22.6015 | -28.1064 | -26.5415 | -13.4405 | 12.5237 | 49.7582 | 94.3511 | 135.2974 | 162.3363 | 167.573 | 151.5308 | 119.0661 | 84.6494 | 54.9722 | 32.4654 | 18.4323 | 8.9305 | 4.1968
Error % | -3.1426 | -2.1641 | -2.787 | -4.5327 | -5.837 | -8.6684 | -16.2977 | -42.6382 | 192.6346 | 34.0453 | 16.2414 | 7.8787 | 3.5309 | 1.1799 | -0.70113 | -2.4195 | -3.8471 | -4.412 | -4.4446 | -4.0259 | -4.4569 | -5.0419 | -5.3498 | -4.8479 | -4.0826 | -4.9494 | -5.9822 | -5.3847 | -2.5588 | 2.2448 | 8.6755 | 17.003 | 27.6219 | 42.9177 | 71.9537 | 177.6029 | -272.2133 | -61.8421 | -28.819 | -15.1828 | -8.6036 | -4.4093 | -2.2832
Sum of Errors = 20.0%

## Slide 17: Row 2 @ 1,200A
Chart: Row #2
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -317.2157 | -361.9076 | -401.6112 | -427.2849 | -424.5877 | -376.4539 | -268.6416 | -104.1075 | 92.8307 | 276.8011 | 413.7939 | 492.6085 | 517.7027 | 508.9125 | 485.4221 | 459.4037 | 435.8437 | 416.6209 | 401.2111 | 362.6981 | 359.7464 | 358.6295 | 359.04 | 361.7287 | 366.708 | 374.3177 | 383.9981 | 394.299 | 401.1791 | 399.0258 | 379.6498 | 336.1415 | 268.2728 | 179.4577 | 83.0622 | -7.9043 | -81.7736 | -132.8812 | -161.6223 | -172.7587 | -171.2342 | -161.5814 | -148.3671
Measured | -307.46 | -355.16 | -392.88 | -411.47 | -398.74 | -339.54 | -218.55 | -42.39 | 157.91 | 339.06 | 465.67 | 528.54 | 541.4 | 522.65 | 491.72 | 459.19 | 431.32 | 410.28 | 393.61 | 355.53 | 351.23 | 349.1 | 349.57 | 353.86 | 356.73 | 360.68 | 369.61 | 384.1 | 399.77 | 412.59 | 414.11 | 394.73 | 349.18 | 277.03 | 185.68 | 89.06 | 1.17 | -69.01 | -115.97 | -142.47 | -151.93 | -150.3 | -141.68
Difference | 9.7557 | 6.7476 | 8.7312 | 15.8149 | 25.8477 | 36.9139 | 50.0916 | 61.7175 | 65.0793 | 62.2589 | 51.8761 | 35.9315 | 23.6973 | 13.7375 | 6.2979 | -0.21365 | -4.5237 | -6.3409 | -7.6011 | -7.1681 | -8.5164 | -9.5295 | -9.47 | -7.8687 | -9.978 | -13.6377 | -14.3881 | -10.199 | -1.4091 | 13.5642 | 34.4602 | 58.5885 | 80.9072 | 97.5723 | 102.6178 | 96.9643 | 82.9436 | 63.8712 | 45.6523 | 30.2887 | 19.3042 | 11.2814 | 6.6871
Error % | -3.173 | -1.8999 | -2.2224 | -3.8435 | -6.4823 | -10.8717 | -22.9199 | -145.5945 | 41.2129 | 18.3622 | 11.1401 | 6.7983 | 4.377 | 2.6284 | 1.2808 | -0.046528 | -1.0488 | -1.5455 | -1.9311 | -2.0162 | -2.4247 | -2.7297 | -2.709 | -2.2237 | -2.7971 | -3.7811 | -3.8928 | -2.6553 | -0.35249 | 3.2876 | 8.3215 | 14.8427 | 23.1706 | 35.2208 | 55.266 | 108.8753 | -92.5535 | -39.3656 | -21.2597 | -12.706 | -7.5059 | -4.7199
Sum of Errors = 15.3%

## Slide 18: Row 3 @ 1,200A
Chart: Row #3
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -247.9799 | -273.4704 | -291.6613 | -294.8788 | -274.6905 | -222.3286 | -135.5328 | -21.2355 | 103.5835 | 216.5296 | 302.8217 | 354.9605 | 376.6198 | 376.6882 | 365.7151 | 350.5632 | 335.2727 | 321.5496 | 310.0242 | 300.8685 | 293.3587 | 287.2945 | 282.3482 | 278.7349 | 275.9441 | 274.7108 | 274.3619 | 275.6257 | 278.3344 | 282.2755 | 286.9045 | 291.3895 | 291.8656 | 285.9708 | 269.6149 | 238.8385 | 194.0564 | 135.7587 | 72.7381 | 11.2738 | -41.1721 | -81.938 | -107.9097 | -121.9514 | -125.8685 | -123.2989 | -116.293 | 6603.7985
Measured | -240.83 | -267.01 | -283.11 | -281.75 | -256.01 | -198.91 | -106.05 | 12.19 | 137.86 | 249.3 | 330.44 | 375.89 | 391.62 | 386.09 | 370.65 | 351.57 | 333.61 | 318.27 | 306.3 | 298.43 | 295.07 | 290.05 | 280.91 | 274.49 | 270.43 | 268.23 | 268.24 | 268.79 | 270.87 | 274.83 | 281.42 | 289.15 | 296.05 | 300.02 | 295.46 | 278.4 | 245.9 | 198.57 | 138.29 | 75.73 | 16.41 | -35.23 | -71.93 | -95.92 | -108.09 | -111.39 | -108.59 | 7298.7
Difference | 7.1499 | 6.4604 | 8.5513 | 13.1288 | 18.6805 | 23.4186 | 29.4828 | 33.4255 | 34.2765 | 32.7704 | 27.6183 | 20.9295 | 15.0002 | 9.4018 | 4.9349 | 1.0068 | -1.6627 | -3.2796 | -3.7242 | -2.4385 | 1.7113 | 2.7555 | -1.4382 | -4.2449 | -5.5141 | -6.4808 | -6.1219 | -6.8357 | -7.4644 | -7.4455 | -5.4845 | -2.2395 | 4.1844 | 14.0492 | 25.8451 | 39.5615 | 51.8436 | 62.8113 | 65.5519 | 64.4562 | 57.5821 | 46.708 | 35.9797 | 26.0314 | 17.7785 | 11.9089 | 7.703 | 694.9015
Error % | -2.9689 | -2.4195 | -3.0205 | -4.6597 | -7.2968 | -11.7735 | -27.8009 | 274.2047 | 24.8633 | 13.1449 | 8.358 | 5.568 | 3.8303 | 2.4351 | 1.3314 | 0.28638 | -0.49841 | -1.0304 | -1.2159 | -0.81711 | 0.57998 | 0.95 | -0.51197 | -1.5465 | -2.039 | -2.4161 | -2.2823 | -2.5432 | -2.7557 | -2.7091 | -1.9489 | -0.77451 | 1.4134 | 4.6828 | 8.7474 | 14.2103 | 21.0832 | 31.6318 | 47.4018 | 85.1132 | 350.8962 | -132.5802 | -50.0205 | -27.1386 | -16.4479 | -10.6911 | -7.0936 | 694.9015
Sum of Errors = 9.5%

## Slide 19: Row 4 @ 1,200A
Chart: Row #4
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -192.4022 | -205.6142 | -211.5552 | -204.8206 | -181.5577 | -137.983 | -75.0397 | 2.0884 | 82.5661 | 155.3108 | 212.0358 | 248.6285 | 267.0242 | 271.6124 | 267.8524 | 259.842 | 250.7445 | 241.8447 | 234.0827 | 227.3912 | 221.7738 | 217.3189 | 213.5049 | 210.6282 | 208.5983 | 207.4072 | 207.0283 | 207.4764 | 208.5514 | 210.1716 | 212.2316 | 212.9001 | 210.8002 | 204.3293 | 190.9436 | 169.2246 | 138.2663 | 99.8226 | 57.1721 | 15.393 | -21.9812 | -51.8916 | -73.4732 | -86.7445 | -92.9605 | -93.8977 | -91.1005 | 4988.2473
Measured | -186.74 | -199.78 | -204.32 | -195.64 | -169.82 | -123.73 | -57.99 | 20.19 | 100.93 | 172.66 | 227.12 | 260.96 | 276.46 | 278.21 | 271.83 | 261.6 | 250.77 | 240.75 | 232.67 | 227.25 | 225.6 | 222.24 | 214.37 | 208.56 | 205.32 | 203.62 | 203.12 | 203.7 | 205.01 | 207.62 | 210.96 | 214.6 | 217.73 | 217.6 | 211.7 | 197.65 | 174.9 | 140.87 | 101.14 | 58.92 | 17.83 | -17.79 | -45.69 | -65.37 | -77.07 | -82.3 | -82.76 | 5482.96
Difference | 5.6622 | 5.8342 | 7.2352 | 9.1806 | 11.7377 | 14.253 | 17.0497 | 18.1016 | 18.3639 | 17.3492 | 15.0842 | 12.3315 | 9.4358 | 6.5976 | 3.9776 | 1.758 | 0.025525 | -1.0947 | -1.4127 | -0.14123 | 3.8262 | 4.9211 | 0.8651 | -2.0682 | -3.2783 | -3.7872 | -3.9083 | -3.7764 | -3.5414 | -2.5516 | -1.2716 | 1.6999 | 6.9298 | 13.2707 | 20.7564 | 28.4254 | 36.6337 | 41.0474 | 43.9679 | 43.527 | 39.8112 | 34.1016 | 27.7832 | 21.3745 | 15.8905 | 11.5977 | 8.3405 | 494.7127
Error % | -3.0321 | -2.9203 | -3.5411 | -4.6926 | -6.9119 | -11.5194 | -29.4011 | 89.6564 | 18.1947 | 10.0482 | 6.6415 | 4.7254 | 3.4131 | 2.3715 | 1.4633 | 0.67203 | 0.010179 | -0.45468 | -0.60715 | -0.062146 | 1.696 | 2.2143 | 0.40356 | -0.99167 | -1.5967 | -1.86 | -1.9241 | -1.8539 | -1.7274 | -1.229 | -0.60276 | 0.79214 | 3.1828 | 6.0987 | 9.8046 | 14.3817 | 20.9455 | 29.1385 | 43.4723 | 73.8748 | 223.2819 | -191.6897 | -60.808 | -32.6977 | -20.6183 | -14.092 | -10.0779 | 494.7127
Sum of Errors = 9.0%

## Slide 20: Row 5 @ 1,200A
Chart: Row #5
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -148.534 | -154.6569 | -154.4586 | -145.2266 | -124.0388 | -90.6299 | -45.7146 | 5.9766 | 59.0842 | 107.1482 | 145.432 | 171.8586 | 187.3247 | 193.7476 | 193.9608 | 190.5885 | 186.0519 | 180.3978 | 175.6531 | 171.025 | 167.0707 | 163.9525 | 161.2611 | 159.053 | 157.3866 | 156.3372 | 155.8504 | 155.8273 | 156.1783 | 156.6597 | 156.7878 | 155.7335 | 152.9104 | 146.5154 | 135.6237 | 119.2707 | 97.866 | 71.4765 | 42.7644 | 13.8111 | -12.4783 | -34.7291 | -51.5298 | -62.9848 | -69.4725 | -72.0116 | -71.458 | 3688.5886
Measured | -143.75 | -149.44 | -148.46 | -138.1 | -115.81 | -81.2 | -35.6 | 16.56 | 69.09 | 116.52 | 153.8 | 178.94 | 192.71 | 197.71 | 196.59 | 191.84 | 186.53 | 180.11 | 174.86 | 171.69 | 169.55 | 166.65 | 162.02 | 158.32 | 156.02 | 154.77 | 154.35 | 154.6 | 155.47 | 156.88 | 158.57 | 160.03 | 160.38 | 158.3 | 151.96 | 139.97 | 122.58 | 99.39 | 72.53 | 43.71 | 15.4 | -9.98 | -30.73 | -46.15 | -56.2 | -61.81 | -63.69 | 4045.33
Difference | 4.784 | 5.2169 | 5.9986 | 7.1266 | 8.2288 | 9.4299 | 10.1146 | 10.5834 | 10.0058 | 9.3718 | 8.368 | 7.0814 | 5.3853 | 3.9624 | 2.6292 | 1.2515 | 0.47815 | -0.28783 | -0.79314 | 0.66503 | 2.4793 | 2.6975 | 0.7589 | -0.73298 | -1.3666 | -1.5672 | -1.5004 | -1.2273 | -0.7083 | 0.22032 | 1.7822 | 4.2965 | 7.4696 | 11.7846 | 16.3363 | 20.6993 | 24.714 | 27.9135 | 29.7656 | 29.8989 | 27.8783 | 24.7491 | 20.7998 | 16.8348 | 13.2725 | 10.2016 | 7.768 | 356.7414
Error % | -3.328 | -3.4909 | -4.0406 | -5.1605 | -7.1054 | -11.6131 | -28.4117 | 63.9094 | 14.4823 | 8.0431 | 5.4408 | 3.9574 | 2.7945 | 2.0041 | 1.3374 | 0.65237 | 0.25634 | -0.15981 | -0.45359 | 0.38734 | 1.4623 | 1.6186 | 0.4684 | -0.46297 | -0.87589 | -1.0126 | -0.97208 | -0.79389 | -0.45559 | 0.14044 | 1.1239 | 2.6848 | 4.6574 | 7.4445 | 10.7504 | 14.7884 | 20.1615 | 28.0848 | 41.039 | 68.4029 | 181.0282 | -247.9875 | -67.6858 | -36.4784 | -23.6165 | -16.5047 | -12.1965 | 356.7414
Sum of Errors = 8.8%

## Slide 21: Row 6 @ 1,200A
Chart: Row #6
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -114.9957 | -117.1932 | -114.3175 | -104.6732 | -87.6693 | -62.538 | -31.1426 | 4.1915 | 39.7011 | 72.4335 | 99.1057 | 118.8054 | 131.2076 | 137.9959 | 140.1691 | 139.4776 | 137.3715 | 134.696 | 131.8286 | 128.9962 | 126.3568 | 124.1507 | 122.185 | 120.6924 | 119.4567 | 118.4672 | 118.1174 | 117.6844 | 117.4608 | 117.1518 | 116.275 | 114.4895 | 111.3797 | 105.7574 | 96.9889 | 84.8618 | 69.3235 | 50.9152 | 30.8805 | 10.8236 | -8.0214 | -24.2833 | -37.2033 | -46.611 | -52.5534 | -55.622 | -56.3011 | 2707.3604
Measured | -111.79 | -113.59 | -110.31 | -100.41 | -82.77 | -57.26 | -25.35 | 10.22 | 45.95 | 78.35 | 104.67 | 123.74 | 135.64 | 141.55 | 143.18 | 141.78 | 139.83 | 135.85 | 132.59 | 129.75 | 127.54 | 125.25 | 122.83 | 120.7 | 119.25 | 118.35 | 118.03 | 118.06 | 118.45 | 119.1 | 119.66 | 119.85 | 118.86 | 115.91 | 110.26 | 101.13 | 88.35 | 72.07 | 52.91 | 32.91 | 12.97 | -5.33 | -20.81 | -32.82 | -41.3 | -46.31 | -49.1 | 2987.92
Difference | 3.2057 | 3.6032 | 4.0075 | 4.2632 | 4.8993 | 5.278 | 5.7926 | 6.0285 | 6.2489 | 5.9165 | 5.5643 | 4.9346 | 4.4324 | 3.5541 | 3.0109 | 2.3024 | 2.4585 | 1.154 | 0.76142 | 0.75384 | 1.1832 | 1.0993 | 0.64502 | 0.0075577 | -0.20673 | -0.11723 | -0.087419 | 0.3756 | 0.98919 | 1.9482 | 3.385 | 5.3605 | 7.4803 | 10.1526 | 13.2711 | 16.2682 | 19.0265 | 21.1548 | 22.0295 | 22.0864 | 20.9914 | 18.9533 | 16.3933 | 13.791 | 11.2534 | 9.312 | 7.2011 | 280.5596
Error % | -2.8676 | -3.1721 | -3.6329 | -4.2458 | -5.9192 | -9.2176 | -22.8505 | 58.9872 | 13.5994 | 7.5514 | 5.3161 | 3.9878 | 3.2678 | 2.5109 | 2.1029 | 1.6239 | 1.7582 | 0.84945 | 0.57426 | 0.581 | 0.9277 | 0.87768 | 0.52513 | 0.0062616 | -0.17336 | -0.09905 | -0.074065 | 0.31814 | 0.83511 | 1.6358 | 2.8289 | 4.4727 | 6.2933 | 8.7591 | 12.0362 | 16.0864 | 21.5354 | 29.3531 | 41.6358 | 67.1115 | 161.846 | -355.5971 | -78.7759 | -42.02 | -27.2479 | -20.108 | -14.6661 | 280.5596
Sum of Errors = 9.4%

## Slide 22: Tosca Field Map @ 4,000A

## Slide 23: Stray By Fields at 4,000A
Chart: HB Stray Field Measurment at 4000A
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Row 6 | -0.24537 | -0.23294 | -0.20131 | -0.14474 | -0.06039 | 0.05253 | 0.18828 | 0.3371 | 0.48521 | 0.62061 | 0.73145 | 0.81297 | 0.8644 | 0.88933 | 0.89427 | 0.88439 | 0.86666 | 0.84571 | 0.82436 | 0.80483 | 0.78901 | 0.77569 | 0.76521 | 0.75848 | 0.75622 | 0.75824 | 0.76393 | 0.77181 | 0.78121 | 0.79092 | 0.79808 | 0.80199 | 0.79567 | 0.77879 | 0.74767 | 0.69953 | 0.63529 | 0.5563 | 0.46615 | 0.37223 | 0.27907 | 0.19274 | 0.11704 | 0.05537 | 0.00757 | -0.02725 | -0.05125
Row 5 | -0.3376 | -0.33406 | -0.30576 | -0.24256 | -0.13668 | 0.01383 | 0.20267 | 0.41468 | 0.63008 | 0.82429 | 0.9795 | 1.0866 | 1.1481 | 1.169 | 1.1625 | 1.1375 | 1.1051 | 1.0699 | 1.037 | 1.0095 | 0.98891 | 0.97202 | 0.95457 | 0.9458 | 0.94026 | 0.94386 | 0.9516 | 0.96382 | 0.97933 | 0.99479 | 1.0096 | 1.0185 | 1.0184 | 1.0028 | 0.96785 | 0.90616 | 0.82058 | 0.7118 | 0.58634 | 0.45575 | 0.32781 | 0.21092 | 0.11205 | 0.03486 | -0.02348 | -0.06288 | -0.08761
Row 4 | -0.46443 | -0.48065 | -0.46538 | -0.40432 | -0.27768 | -0.08011 | 0.18681 | 0.49869 | 0.81887 | 1.1067 | 1.3288 | 1.4722 | 1.5417 | 1.5526 | 1.5241 | 1.4761 | 1.4203 | 1.3648 | 1.3162 | 1.2773 | 1.2527 | 1.23 | 1.2032 | 1.1874 | 1.1831 | 1.1868 | 1.1986 | 1.2169 | 1.2389 | 1.265 | 1.2896 | 1.3096 | 1.3211 | 1.3115 | 1.2722 | 1.1956 | 1.0796 | 0.92381 | 0.74772 | 0.55857 | 0.37887 | 0.2194 | 0.09052 | -0.00726 | -0.07505 | -0.11627 | -0.13897
Row 3 | -0.62799 | -0.68096 | -0.69873 | -0.65617 | -0.52404 | -0.27901 | 0.09103 | 0.58137 | 1.0372 | 1.4731 | 1.7986 | 1.9896 | 2.0647 | 2.0514 | 1.9882 | 1.9017 | 1.8129 | 1.7312 | 1.6599 | 1.607 | 1.5719 | 1.5416 | 1.5102 | 1.4917 | 1.4865 | 1.4925 | 1.5092 | 1.5327 | 1.5667 | 1.6057 | 1.647 | 1.6904 | 1.7159 | 1.718 | 1.678 | 1.5809 | 1.4239 | 1.2089 | 0.94278 | 0.67999 | 0.43037 | 0.20054 | 0.03765 | -0.0861 | -0.15422 | -0.19374 | -0.20933
Row 2 | -0.83319 | -0.94798 | -1.0297 | -1.0477 | -0.95057 | -0.68767 | -0.20722 | 0.46188 | 1.2165 | 1.9011 | 2.3925 | 2.653 | 2.7224 | 2.6621 | 2.5406 | 2.3993 | 2.2655 | 2.1547 | 2.0597 | 1.8544 | 1.8518 | 1.8619 | 1.8867 | 1.9254 | 1.9695 | 2.0241 | 2.089 | 2.1614 | 2.2196 | 2.2489 | 2.2187 | 2.0954 | 1.8827 | 1.5704 | 1.1887 | 0.79449 | 0.43089 | 0.13546 | -0.07516 | -0.2093 | -0.28016 | -0.30827 | -0.30728
Row 1 | -1.0787 | -1.2774 | -1.4594 | -1.6017 | -1.6515 | -1.5003 | -1.0191 | -0.13966 | 1.0274 | 2.163 | 2.9659 | 3.343 | 3.3878 | 3.2471 | 3.0413 | 2.8338 | 2.6539 | 2.5139 | 2.4082 | 2.2066 | 2.2113 | 2.2302 | 2.2664 | 2.3231 | 2.4005 | 2.4846 | 2.5827 | 2.6998 | 2.8183 | 2.9059 | 2.9112 | 2.7822 | 2.4795 | 2.0204 | 1.42 | 0.85651 | 0.33962 | -0.04174 | -0.28564 | -0.41483 | -0.46374 | -0.46246 | -0.43314

## Slide 24: Row 1 @ 4,000A
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -1135.3154 | -1339.3489 | -1539.6242 | -1714.8398 | -1797.5915 | -1702.0389 | -1315.4882 | -540.3358 | 538.0472 | 1667.1095 | 2557.9736 | 3064.5755 | 3198.3099 | 3106.0815 | 2936.6213 | 2756.5765 | 2598.5074 | 2465.7977 | 2362.6386 | 2167.7836 | 2182.0407 | 2213.4866 | 2260.7981 | 2325.8892 | 2408.7803 | 2506.7859 | 2614.3813 | 2719.2926 | 2797.1523 | 2808.9225 | 2698.9823 | 2425.6738 | 1992.2032 | 1440.0588 | 845.5859 | 318.8031 | -83.2176 | -349.5273 | -490.9633 | -544.5671 | -542.1308 | -507.5931 | -459.309
Measured | -1078.67 | -1277.36 | -1459.45 | -1601.65 | -1651.53 | -1500.31 | -1019.14 | -139.66 | 1027.43 | 2163 | 2965.89 | 3342.97 | 3387.82 | 3247.13 | 3041.27 | 2833.76 | 2653.91 | 2513.85 | 2408.23 | 2206.57 | 2211.29 | 2230.23 | 2266.41 | 2323.07 | 2400.54 | 2484.62 | 2582.66 | 2699.84 | 2818.29 | 2905.88 | 2911.2 | 2782.24 | 2479.49 | 2020.42 | 1419.99 | 856.51 | 339.62 | -41.74 | -285.64 | -414.83 | -463.74 | -462.46 | -433.14
Difference | 56.6454 | 61.9889 | 80.1742 | 113.1898 | 146.0615 | 201.7289 | 296.3482 | 400.6758 | 489.3828 | 495.8905 | 407.9164 | 278.3945 | 189.5101 | 141.0485 | 104.6487 | 77.1835 | 55.4026 | 48.0523 | 45.5914 | 38.7864 | 29.2493 | 16.7434 | 5.6119 | -2.8192 | -8.2403 | -22.1659 | -31.7213 | -19.4526 | 21.1377 | 96.9575 | 212.2177 | 356.5662 | 487.2868 | 580.3612 | 574.4041 | 537.7069 | 422.8376 | 307.7873 | 205.3233 | 129.7371 | 78.3908 | 45.1331 | 26.169
Error % | -5.2514 | -4.8529 | -5.4935 | -7.0671 | -8.844 | -13.4458 | -29.0783 | -286.8937 | 47.6317 | 22.9261 | 13.7536 | 8.3278 | 5.5939 | 4.3438 | 3.441 | 2.7237 | 2.0876 | 1.9115 | 1.8931 | 1.7578 | 1.3227 | 0.75075 | 0.24761 | -0.12135 | -0.34327 | -0.89213 | -1.2282 | -0.72051 | 0.75002 | 3.3366 | 7.2897 | 12.8158 | 19.6527 | 28.7248 | 40.4513 | 62.7788 | 124.5032 | -737.3918 | -71.8819 | -31.2748 | -16.904 | -9.7594 | -6.0417
Sum of Errors = 11.8%

## Slide 25: Row 2 @ 4,000A
Chart: Row #2
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -879.8942 | -995.6005 | -1088.5016 | -1126.961 | -1063.4628 | -839.1154 | -409.0266 | 214.868 | 947.3043 | 1633.7004 | 2151.4116 | 2455.417 | 2558.7491 | 2526.6503 | 2426.7239 | 2305.8486 | 2186.8747 | 2082.1048 | 1994.0406 | 1802.2354 | 1805.9744 | 1824.1581 | 1854.2571 | 1897.5592 | 1951.3644 | 2013.7612 | 2078.5151 | 2135.6983 | 2166.8115 | 2149.4794 | 2055.2084 | 1861.7936 | 1574.0935 | 1209.2925 | 814.5552 | 441.0784 | 129.6929 | -98.66 | -245.2847 | -325.952 | -358.1204 | -357.1453 | -339.6077
Measured | -833.19 | -947.98 | -1029.67 | -1047.72 | -950.57 | -687.67 | -207.22 | 461.88 | 1216.51 | 1901.13 | 2392.48 | 2653.01 | 2722.41 | 2662.07 | 2540.59 | 2399.26 | 2265.49 | 2154.71 | 2059.72 | 1854.4 | 1851.77 | 1886.72 | 1925.4 | 1969.5 | 2024.07 | 2088.99 | 2161.35 | 2219.59 | 2248.92 | 2218.73 | 2095.41 | 1882.68 | 1570.42 | 1188.68 | 794.49 | 430.89 | 135.46 | -75.16 | -209.3 | -280.16 | -308.27 | -307.28
Difference | 46.7042 | 47.6205 | 58.8316 | 79.241 | 112.8928 | 151.4454 | 201.8066 | 247.012 | 269.2057 | 267.4296 | 241.0684 | 197.593 | 163.6609 | 135.4197 | 113.8661 | 93.4114 | 78.6153 | 72.6052 | 65.6794 | 52.1646 | 45.7956 | 32.4629 | 27.8408 | 18.1356 | 10.3088 | 10.4749 | 25.6517 | 52.7785 | 99.4406 | 163.5216 | 233.6164 | 308.5865 | 361.1275 | 374.1248 | 353.4116 | 301.1971 | 234.12 | 170.1247 | 116.652 | 77.9604 | 48.8753 | 32.3277
Error % | -5.6055 | -5.0234 | -5.7136 | -7.5632 | -11.8763 | -22.023 | -97.3876 | 53.4797 | 22.1293 | 14.0669 | 10.0761 | 7.4479 | 6.0116 | 5.087 | 4.4819 | 3.8933 | 3.4701 | 3.3696 | 3.1888 | 2.813 | 2.4731 | 1.7206 | 1.446 | 0.92082 | 0.50931 | 0.50143 | 1.1868 | 2.3778 | 4.4217 | 7.3701 | 11.149 | 16.3908 | 22.9956 | 31.474 | 44.4828 | 69.9012 | 172.8333 | -226.35 | -55.7344 | -27.8271 | -15.8547 | -10.5206
Sum of Errors = 13.5%

## Slide 26: Row 3 @ 4,000A
Chart: Row #3
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -664.3959 | -721.2738 | -747.8291 | -719.8546 | -606.1425 | -379.8914 | -34.4117 | 406.3292 | 882.2463 | 1315.1183 | 1648.4705 | 1855.3051 | 1944.1391 | 1944.2459 | 1893.8854 | 1819.5534 | 1739.6231 | 1663.5595 | 1597.4581 | 1543.9781 | 1501.8062 | 1471.2534 | 1451.3023 | 1442.8092 | 1443.3499 | 1455.3163 | 1475.341 | 1504.9125 | 1540.924 | 1580.4265 | 1618.1967 | 1648.5715 | 1654.0028 | 1624.8004 | 1547.1562 | 1407.8697 | 1212.1224 | 963.9773 | 697.8594 | 438.3907 | 211.9112 | 29.1508 | -99.2615 | -182.5579 | -227.2276 | -245.5926 | -245.1502 | 44202.3018
Measured | -627.99 | -680.96 | -698.73 | -656.17 | -524.04 | -279.01 | 91.03 | 581.37 | 1037.24 | 1473.06 | 1798.64 | 1989.65 | 2064.73 | 2051.37 | 1988.22 | 1901.69 | 1812.88 | 1731.16 | 1659.92 | 1607.03 | 1571.93 | 1541.62 | 1510.16 | 1491.74 | 1486.47 | 1492.53 | 1509.17 | 1532.65 | 1566.65 | 1605.66 | 1646.96 | 1690.38 | 1715.91 | 1717.95 | 1677.98 | 1580.85 | 1423.91 | 1208.91 | 942.78 | 679.99 | 430.37 | 200.54 | 37.65 | -86.1 | -154.22 | -193.74 | -209.33 | 48583.85
Difference | 36.4059 | 40.3138 | 49.0991 | 63.6846 | 82.1025 | 100.8814 | 125.4417 | 175.0408 | 154.9937 | 157.9417 | 150.1695 | 134.3449 | 120.5909 | 107.1241 | 94.3346 | 82.1366 | 73.2569 | 67.6005 | 62.4619 | 63.0519 | 70.1238 | 70.3666 | 58.8577 | 48.9308 | 43.1201 | 37.2137 | 33.829 | 27.7375 | 25.726 | 25.2335 | 28.7633 | 41.8085 | 61.9072 | 93.1496 | 130.8238 | 172.9803 | 211.7876 | 244.9327 | 244.9206 | 241.5993 | 218.4588 | 171.3892 | 136.9115 | 96.4579 | 73.0076 | 51.8526 | 35.8202 | 4381.5482
Error % | -5.7972 | -5.9201 | -7.0269 | -9.7055 | -15.6672 | -36.1569 | 137.8026 | 30.1083 | 14.9429 | 10.722 | 8.3491 | 6.7522 | 5.8405 | 5.2221 | 4.7447 | 4.3191 | 4.0409 | 3.9049 | 3.7629 | 3.9235 | 4.461 | 4.5645 | 3.8974 | 3.2801 | 2.9008 | 2.4933 | 2.2416 | 1.8098 | 1.6421 | 1.5715 | 1.7464 | 2.4733 | 3.6078 | 5.4221 | 7.7965 | 10.9422 | 14.8737 | 20.2606 | 25.9786 | 35.5298 | 50.7607 | 85.4639 | 363.6427 | -112.0301 | -47.3399 | -26.764 | -17.1118 | 4381.5482
Sum of Errors = 9.0%

## Slide 27: Row 4 @ 4,000A
Chart: Row #4
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -494.0405 | -513.7918 | -505.3287 | -450.238 | -335.6526 | -149.143 | 105.1368 | 408.6515 | 722.8594 | 1007.5546 | 1232.318 | 1379.7745 | 1454.8908 | 1472.0259 | 1451.0625 | 1408.3044 | 1357.7343 | 1306.5095 | 1260.431 | 1220.9293 | 1189.2127 | 1166.3741 | 1150.6861 | 1143.2882 | 1143.8106 | 1151.538 | 1165.7508 | 1185.3051 | 1207.9276 | 1231.3633 | 1253.7544 | 1265.1875 | 1259.3718 | 1229.7593 | 1167.1303 | 1067.3697 | 928.923 | 760.2795 | 575.6208 | 393.1896 | 227.8152 | 90.4359 | -15.3592 | -89.8017 | -136.932 | -163.0261 | -173.4303 | 35278.7222
Measured | -464.43 | -480.65 | -465.38 | -404.32 | -277.68 | -80.11 | 186.81 | 498.69 | 818.87 | 1106.69 | 1328.81 | 1472.21 | 1541.73 | 1552.57 | 1524.12 | 1476.06 | 1420.29 | 1364.84 | 1316.18 | 1277.28 | 1252.66 | 1230.01 | 1203.24 | 1187.43 | 1183.11 | 1186.77 | 1198.64 | 1216.88 | 1238.95 | 1265.01 | 1289.62 | 1309.61 | 1321.09 | 1311.51 | 1272.17 | 1195.57 | 1079.62 | 923.81 | 747.72 | 558.57 | 378.87 | 219.4 | 90.52 | -7.26 | -75.05 | -116.27 | -138.97 | 38573.36
Difference | 29.6105 | 33.1418 | 39.9487 | 45.918 | 57.9726 | 69.033 | 81.6732 | 90.0385 | 96.0106 | 99.1354 | 96.492 | 92.4355 | 86.8392 | 80.5441 | 73.0575 | 67.7556 | 62.5557 | 58.3305 | 55.749 | 56.3507 | 63.4473 | 63.6359 | 52.5539 | 44.1418 | 39.2994 | 35.232 | 32.8892 | 31.5749 | 31.0224 | 33.6467 | 35.8656 | 44.4225 | 61.7182 | 81.7507 | 105.0397 | 128.2003 | 150.697 | 163.5305 | 172.0992 | 165.3804 | 151.0548 | 128.9641 | 105.8792 | 82.5417 | 61.882 | 46.7561 | 34.4603 | 3294.6378
Error % | -6.3757 | -6.8952 | -8.5841 | -11.3568 | -20.8775 | -86.1728 | 43.7199 | 18.055 | 11.7248 | 8.9578 | 7.2615 | 6.2787 | 5.6326 | 5.1878 | 4.7934 | 4.5903 | 4.4044 | 4.2738 | 4.2357 | 4.4118 | 5.065 | 5.1736 | 4.3677 | 3.7174 | 3.3217 | 2.9687 | 2.7439 | 2.5947 | 2.5039 | 2.6598 | 2.7811 | 3.392 | 4.6718 | 6.2333 | 8.2567 | 10.7229 | 13.9583 | 17.7017 | 23.0165 | 29.6078 | 39.8698 | 58.7804 | 116.9677 | -1136.9375 | -82.4544 | -40.2134 | -24.7969 | 3294.6378
Sum of Errors = 8.5%

## Slide 28: Row 5 @ 4,000A
Chart: Row #5
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -362.2841 | -362.1843 | -337.9076 | -280.4364 | -179.9096 | -35.8723 | 148.7432 | 356.6693 | 568.0976 | 760.8672 | 915.22 | 1022.8006 | 1085.8983 | 1110.0462 | 1106.1276 | 1084.4412 | 1056.106 | 1022.002 | 991.9848 | 963.9473 | 941.2341 | 924.8328 | 913.6009 | 907.5633 | 907.0555 | 911.8879 | 921.2406 | 934.0358 | 948.4212 | 962.4711 | 972.7804 | 974.6408 | 965.2457 | 936.8191 | 886.2601 | 810.717 | 713.1123 | 594.5139 | 467.0279 | 336.7426 | 217.4147 | 113.1988 | 29.6543 | -33.2583 | -76.9058 | -104.7402 | -119.7967 | 27924.8281
Measured | -337.6 | -334.06 | -305.76 | -242.56 | -136.68 | 13.83 | 202.67 | 414.68 | 630.08 | 824.29 | 979.5 | 1086.57 | 1148.09 | 1169.03 | 1162.46 | 1137.49 | 1105.06 | 1069.92 | 1036.99 | 1009.49 | 988.91 | 972.02 | 954.57 | 945.8 | 940.26 | 943.86 | 951.6 | 963.82 | 979.33 | 994.79 | 1009.59 | 1018.46 | 1018.41 | 1002.83 | 967.85 | 906.16 | 820.58 | 711.8 | 586.34 | 455.75 | 327.81 | 210.92 | 112.05 | 34.86 | -23.48 | -62.88 | -87.61 | 30417
Difference | 24.6841 | 28.1243 | 32.1476 | 37.8764 | 43.2296 | 49.7023 | 53.9268 | 58.0107 | 61.9824 | 63.4228 | 64.28 | 63.7694 | 62.1917 | 58.9838 | 56.3324 | 53.0488 | 48.954 | 47.918 | 45.0052 | 45.5427 | 47.6759 | 47.1872 | 40.9691 | 38.2367 | 33.2045 | 31.9721 | 30.3594 | 29.7842 | 30.9088 | 32.3189 | 36.8096 | 43.8192 | 53.1643 | 66.0109 | 81.5899 | 95.443 | 107.4677 | 117.2861 | 119.3121 | 119.0074 | 110.3953 | 97.7212 | 82.3957 | 68.1183 | 53.4258 | 41.8602 | 32.1867 | 2492.1719
Error % | -7.3116 | -8.4189 | -10.514 | -15.6153 | -31.6283 | 359.38 | 26.6082 | 13.9893 | 9.8372 | 7.6942 | 6.5625 | 5.8689 | 5.417 | 5.0455 | 4.846 | 4.6637 | 4.43 | 4.4787 | 4.34 | 4.5115 | 4.8211 | 4.8546 | 4.2919 | 4.0428 | 3.5314 | 3.3874 | 3.1904 | 3.0902 | 3.1561 | 3.2488 | 3.646 | 4.3025 | 5.2203 | 6.5825 | 8.43 | 10.5327 | 13.0966 | 16.4774 | 20.3486 | 26.1124 | 33.6766 | 46.3309 | 73.5348 | 195.4053 | -227.5376 | -66.5716 | -36.7386 | 2492.1719
Sum of Errors = 8.2%

## Slide 29: Row 6 @ 4,000A
Chart: Row #6
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -263.8771 | -254.0963 | -225.0153 | -171.1451 | -90.5442 | 19.1381 | 150.7299 | 295.8829 | 440.601 | 574.0629 | 683.1093 | 763.9289 | 814.3114 | 840.3624 | 845.4347 | 836.7689 | 820.8599 | 802.014 | 782.6019 | 764.3715 | 748.6473 | 736.9143 | 728.5365 | 724.6203 | 723.9569 | 726.5106 | 733.2523 | 740.3517 | 748.6314 | 755.771 | 758.9561 | 755.9396 | 744.7216 | 720.052 | 679.5798 | 623.1855 | 551.3459 | 466.8696 | 375.801 | 283.1937 | 195.4326 | 117.4537 | 52.253 | 0.64885 | -37.3567 | -63.7804 | -80.4147 | 22121.4759
Measured | -245.37 | -232.94 | -201.31 | -144.74 | -60.39 | 52.53 | 188.28 | 337.1 | 485.21 | 620.61 | 731.45 | 812.97 | 864.4 | 889.33 | 894.27 | 884.39 | 866.66 | 845.71 | 824.36 | 804.83 | 789.01 | 775.69 | 765.21 | 758.48 | 756.22 | 758.24 | 763.93 | 771.81 | 781.21 | 790.92 | 798.08 | 801.99 | 795.67 | 778.79 | 747.67 | 699.53 | 635.29 | 556.3 | 466.15 | 372.23 | 279.07 | 192.74 | 117.04 | 55.37 | 7.57 | -27.25 | 24168.62
Difference | 18.5071 | 21.1563 | 23.7053 | 26.4051 | 30.1542 | 33.3919 | 37.5501 | 41.2171 | 44.609 | 46.5471 | 48.3407 | 49.0411 | 50.0886 | 48.9676 | 48.8353 | 47.6211 | 45.8001 | 43.696 | 41.7581 | 40.4585 | 40.3627 | 38.7757 | 36.6735 | 33.8597 | 32.2631 | 31.7294 | 30.6777 | 31.4583 | 32.5786 | 35.149 | 39.1239 | 46.0504 | 50.9484 | 58.738 | 68.0902 | 76.3445 | 83.9441 | 89.4304 | 90.349 | 89.0363 | 83.6374 | 75.2863 | 64.787 | 54.7211 | 44.9267 | 36.5304 | 2047.1441
Error % | -7.5425 | -9.0823 | -11.7755 | -18.2431 | -49.9324 | 63.5674 | 19.9438 | 12.227 | 9.1937 | 7.5002 | 6.6089 | 6.0323 | 5.7946 | 5.5061 | 5.4609 | 5.3846 | 5.2847 | 5.1668 | 5.0655 | 5.027 | 5.1156 | 4.9989 | 4.7926 | 4.4642 | 4.2664 | 4.1846 | 4.0158 | 4.0759 | 4.1703 | 4.4441 | 4.9023 | 5.742 | 6.4032 | 7.5422 | 9.107 | 10.9137 | 13.2135 | 16.0759 | 19.382 | 23.9197 | 29.97 | 39.0611 | 55.3546 | 98.8282 | 593.483 | -134.0564 | 2047.1441
Sum of Errors = 8.5%

## Slide 30: Coil translations in Tosca
Moved the coil in X direction -3 mm and +3mm.
A +3mm matched the measured data better at higher excitations.
X direction
- +

## Slide 31: Row 1 @ 1,200A Coils moved +3 mm in X
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca +3mm | -407.5309 | -483.9077 | -562.9573 | -639.3371 | -693.3002 | -698.9708 | -619.4442 | -417.6432 | -108.7925 | 222.719 | 492.5549 | 645.3792 | 681.2062 | 654.0756 | 607.4364 | 567.6088 | 536.7707 | 512.3431 | 494.8305 | 462.4283 | 461.7378 | 463.1558 | 466.5753 | 473.0908 | 483.9193 | 499.3158 | 519.6322 | 543.1729 | 564.712 | 573.223 | 551.7494 | 485.3081 | 372.4943 | 224.791 | 64.6969 | -73.9862 | -174.4214 | -234.2505 | -257.6305 | -256.5634 | -241.1532 | -218.2407 | -193.351
Measured | -386.15 | -460.93 | -529.98 | -587.31 | -622.13 | -600.7 | -483.96 | -251.37 | 66.99 | 374.89 | 587.59 | 681.45 | 685.33 | 643.52 | 589.7 | 541.32 | 503.94 | 478.38 | 461.8 | 429.77 | 426.83 | 425.2 | 426.62 | 434.2 | 446.83 | 456.65 | 469.83 | 492.91 | 525.26 | 557.91 | 573.55 | 554.91 | 489.82 | 378.25 | 232.89 | 85.32 | -43.74 | -136.88 | -190.75 | -213.83 | -214.24 | -202.54 | -183.81
Tosca 0 | -398.2851 | -470.905 | -544.7506 | -613.9313 | -658.4438 | -652.7708 | -562.8342 | -358.5496 | -62.0559 | 247.2577 | 492.1569 | 627.7608 | 661.1316 | 635.9268 | 593.8346 | 554.4172 | 523.327 | 499.4859 | 482.3252 | 447.072 | 445.8535 | 446.6381 | 449.4433 | 455.2496 | 465.0723 | 479.2515 | 497.9364 | 519.4515 | 538.7005 | 545.3863 | 523.7918 | 460.5589 | 354.5226 | 215.9137 | 65.317 | -66.2108 | -162.8061 | -221.5294 | -245.7222 | -246.2954 | -232.6723 | -211.4705 | -188.0068
Difference | 21.3809 | 22.9777 | 32.9773 | 52.0271 | 71.1702 | 98.2708 | 135.4842 | 166.2732 | 175.7825 | 152.171 | 95.0351 | 36.0708 | 4.1238 | -10.5556 | -17.7364 | -26.2888 | -32.8307 | -33.9631 | -33.0305 | -32.6583 | -34.9078 | -37.9558 | -39.9553 | -38.8908 | -37.0893 | -42.6658 | -49.8022 | -50.2629 | -39.452 | -15.313 | 21.8006 | 69.6019 | 117.3257 | 153.459 | 168.1931 | 159.3062 | 130.6814 | 97.3705 | 66.8805 | 42.7334 | 26.9132 | 15.7007 | 9.541
Error % | -5.5369 | -4.9851 | -6.2224 | -8.8585 | -11.4398 | -16.3594 | -27.9949 | -66.1468 | 262.401 | 40.5908 | 16.1737 | 5.2932 | 0.60172 | -1.6403 | -3.0077 | -4.8564 | -6.5148 | -7.0996 | -7.1525 | -7.599 | -8.1784 | -8.9266 | -9.3656 | -8.9569 | -8.3005 | -9.3432 | -10.6 | -10.1972 | -7.5109 | -2.7447 | 3.801 | 12.5429 | 23.9528 | 40.5708 | 72.22 | 186.7162 | -298.7686 | -71.1357 | -35.0619 | -19.9847 | -12.5622 | -7.7519 | -5.1907
Sum of Errors = 19.8%

## Slide 32: Row 6 @ 1,200A Coils moved +3 mm in X
Chart: Row #6
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca +3mm | -118.1528 | -120.7025 | -118.0063 | -108.2322 | -90.6717 | -64.4337 | -31.4159 | 5.9349 | 43.5968 | 78.3702 | 106.7186 | 127.6182 | 140.7158 | 147.8133 | 149.9958 | 149.1451 | 146.8086 | 143.8893 | 140.7867 | 137.7394 | 134.9112 | 132.5516 | 130.4552 | 128.8669 | 127.558 | 126.5196 | 126.1646 | 125.7391 | 125.5503 | 125.2899 | 124.452 | 122.67 | 119.4954 | 113.6637 | 104.4809 | 91.7096 | 75.2854 | 55.7852 | 34.5619 | 13.2804 | -6.6772 | -23.8675 | -37.4889 | -47.3802 | -53.6037 | -56.8017 | -57.4976 | 2938.4749
Measured | -111.79 | -113.59 | -110.31 | -100.41 | -82.77 | -57.26 | -25.35 | 10.22 | 45.95 | 78.35 | 104.67 | 123.74 | 135.64 | 141.55 | 143.18 | 141.78 | 139.83 | 135.85 | 132.59 | 129.75 | 127.54 | 125.25 | 122.83 | 120.7 | 119.25 | 118.35 | 118.03 | 118.06 | 118.45 | 119.1 | 119.66 | 119.85 | 118.86 | 115.91 | 110.26 | 101.13 | 88.35 | 72.07 | 52.91 | 32.91 | 12.97 | -5.33 | -20.81 | -32.82 | -41.3 | -46.31 | 2987.92
Tosca 0 | -114.9957 | -117.1932 | -114.3175 | -104.6732 | -87.6693 | -62.538 | -31.1426 | 4.1915 | 39.7011 | 72.4335 | 99.1057 | 118.8054 | 131.2076 | 137.9959 | 140.1691 | 139.4776 | 137.3715 | 134.696 | 131.8286 | 128.9962 | 126.3568 | 124.1507 | 122.185 | 120.6924 | 119.4567 | 118.4672 | 118.1174 | 117.6844 | 117.4608 | 117.1518 | 116.275 | 114.4895 | 111.3797 | 105.7574 | 96.9889 | 84.8618 | 69.3235 | 50.9152 | 30.8805 | 10.8236 | -8.0214 | -24.2833 | -37.2033 | -46.611 | -52.5534 | -55.622 | -56.3011
Difference | 6.3628 | 7.1125 | 7.6963 | 7.8222 | 7.9017 | 7.1737 | 6.0659 | 4.2851 | 2.3532 | -0.02016 | -2.0486 | -3.8782 | -5.0758 | -6.2633 | -6.8158 | -7.3651 | -6.9786 | -8.0393 | -8.1967 | -7.9894 | -7.3712 | -7.3016 | -7.6252 | -8.1669 | -8.308 | -8.1696 | -8.1346 | -7.6791 | -7.1003 | -6.1899 | -4.792 | -2.82 | -0.63537 | 2.2463 | 5.7791 | 9.4204 | 13.0646 | 16.2848 | 18.3481 | 19.6296 | 19.6472 | 18.5375 | 16.6789 | 14.5602 | 12.3037 | 10.4917 | 49.4451
Error % | -5.6917 | -6.2616 | -6.977 | -7.7903 | -9.5466 | -12.5282 | -23.9285 | 41.9282 | 5.1212 | -0.02573 | -1.9572 | -3.1342 | -3.7421 | -4.4248 | -4.7603 | -5.1947 | -4.9908 | -5.9177 | -6.182 | -6.1575 | -5.7795 | -5.8296 | -6.2079 | -6.7663 | -6.9669 | -6.9029 | -6.892 | -6.5044 | -5.9943 | -5.1972 | -4.0047 | -2.3529 | -0.53455 | 1.9379 | 5.2414 | 9.3151 | 14.7874 | 22.5958 | 34.6779 | 59.6462 | 151.4818 | -347.7949 | -80.1483 | -44.3637 | -29.7911 | -22.6554 | 49.4451
Sum of Errors = 1.7%

## Slide 33: Row 1 @ 4,000A Coils moved +3 mm in X
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca +3mm | -1159.2691 | -1373.6543 | -1588.3662 | -1783.6363 | -1892.6693 | -1828.4117 | -1469.6018 | -697.528 | 428.8221 | 1623.8159 | 2594.1606 | 3149.191 | 3299.3804 | 3203.0005 | 3025.2782 | 2840.2637 | 2678.9289 | 2543.5787 | 2438.092 | 2247.0907 | 2263.5105 | 2298.2171 | 2348.1969 | 2416.842 | 2504.28 | 2607.3446 | 2720.8981 | 2832.4043 | 2916.7783 | 2932.6487 | 2820.6008 | 2534.2495 | 2076.3845 | 1492.6282 | 865.4164 | 313.6806 | -103.5223 | -376.004 | -517.2717 | -567.8073 | -561.4129 | -522.8875 | -471.2302
Measured | -1078.67 | -1277.36 | -1459.45 | -1601.65 | -1651.53 | -1500.31 | -1019.14 | -139.66 | 1027.43 | 2163 | 2965.89 | 3342.97 | 3387.82 | 3247.13 | 3041.27 | 2833.76 | 2653.91 | 2513.85 | 2408.23 | 2206.57 | 2211.29 | 2230.23 | 2266.41 | 2323.07 | 2400.54 | 2484.62 | 2582.66 | 2699.84 | 2818.29 | 2905.88 | 2911.2 | 2782.24 | 2479.49 | 2020.42 | 1419.99 | 856.51 | 339.62 | -41.74 | -285.64 | -414.83 | -463.74 | -462.46 | -433.14
Tosca 0 | -1135.3154 | -1339.3489 | -1539.6242 | -1714.8398 | -1797.5915 | -1702.0389 | -1315.4882 | -540.3358 | 538.0472 | 1667.1095 | 2557.9736 | 3064.5755 | 3198.3099 | 3106.0815 | 2936.6213 | 2756.5765 | 2598.5074 | 2465.7977 | 2362.6386 | 2167.7836 | 2182.0407 | 2213.4866 | 2260.7981 | 2325.8892 | 2408.7803 | 2506.7859 | 2614.3813 | 2719.2926 | 2797.1523 | 2808.9225 | 2698.9823 | 2425.6738 | 1992.2032 | 1440.0588 | 845.5859 | 318.8031 | -83.2176 | -349.5273 | -490.9633 | -544.5671 | -542.1308 | -507.5931 | -459.309
Difference | 80.5991 | 96.2943 | 128.9162 | 181.9863 | 241.1393 | 328.1017 | 450.4618 | 557.868 | 598.6079 | 539.1841 | 371.7294 | 193.779 | 88.4396 | 44.1295 | 15.9918 | -6.5037 | -25.0189 | -29.7287 | -29.862 | -40.5207 | -52.2205 | -67.9871 | -81.7869 | -93.772 | -103.74 | -122.7246 | -138.2381 | -132.5643 | -98.4883 | -26.7687 | 90.5992 | 247.9905 | 403.1055 | 527.7918 | 554.5736 | 542.8294 | 443.1423 | 334.264 | 231.6317 | 152.9773 | 97.6729 | 60.4275 | 38.0902
Error % | -7.4721 | -7.5385 | -8.8332 | -11.3624 | -14.601 | -21.8689 | -44.2002 | -399.4472 | 58.2626 | 24.9276 | 12.5335 | 5.7966 | 2.6105 | 1.359 | 0.52583 | -0.22951 | -0.94272 | -1.1826 | -1.24 | -1.8364 | -2.3615 | -3.0484 | -3.6087 | -4.0366 | -4.3215 | -4.9394 | -5.3525 | -4.9101 | -3.4946 | -0.92119 | 3.1121 | 8.9133 | 16.2576 | 26.1229 | 39.0548 | 63.3769 | 130.4818 | -800.8242 | -81.0922 | -36.8771 | -21.062 | -13.0665 | -8.794
Sum of Errors = 11.4%

## Slide 34: Row 6 @ 4,000A Coils moved +3 mm in X
Chart: Row #6
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca +3mm | -270.1177 | -260.7518 | -231.4476 | -176.2472 | -92.8956 | 21.3298 | 159.1889 | 311.8771 | 464.5592 | 605.6246 | 721.026 | 806.551 | 859.8156 | 887.2822 | 892.5469 | 883.3005 | 866.4115 | 846.4361 | 825.8755 | 806.5716 | 789.914 | 777.4586 | 768.5364 | 764.3093 | 763.5066 | 766.0942 | 773.086 | 780.4791 | 789.1345 | 796.6288 | 800.0198 | 796.9355 | 785.2562 | 759.447 | 716.9805 | 657.6953 | 582.0762 | 493.0954 | 397.1661 | 299.6738 | 207.3934 | 125.544 | 57.2592 | 3.3531 | -36.2369 | -63.6835 | -80.91 | 23374.6267
Measured | -245.37 | -232.94 | -201.31 | -144.74 | -60.39 | 52.53 | 188.28 | 337.1 | 485.21 | 620.61 | 731.45 | 812.97 | 864.4 | 889.33 | 894.27 | 884.39 | 866.66 | 845.71 | 824.36 | 804.83 | 789.01 | 775.69 | 765.21 | 758.48 | 756.22 | 758.24 | 763.93 | 771.81 | 781.21 | 790.92 | 798.08 | 801.99 | 795.67 | 778.79 | 747.67 | 699.53 | 635.29 | 556.3 | 466.15 | 372.23 | 279.07 | 192.74 | 117.04 | 55.37 | 7.57 | -27.25 | 24168.62
Tosca 0 | -263.8771 | -254.0963 | -225.0153 | -171.1451 | -90.5442 | 19.1381 | 150.7299 | 295.8829 | 440.601 | 574.0629 | 683.1093 | 763.9289 | 814.3114 | 840.3624 | 845.4347 | 836.7689 | 820.8599 | 802.014 | 782.6019 | 764.3715 | 748.6473 | 736.9143 | 728.5365 | 724.6203 | 723.9569 | 726.5106 | 733.2523 | 740.3517 | 748.6314 | 755.771 | 758.9561 | 755.9396 | 744.7216 | 720.052 | 679.5798 | 623.1855 | 551.3459 | 466.8696 | 375.801 | 283.1937 | 195.4326 | 117.4537 | 52.253 | 0.64885 | -37.3567 | -63.7804 | -80.4147
Difference | 24.7477 | 27.8118 | 30.1376 | 31.5072 | 32.5056 | 31.2002 | 29.0911 | 25.2229 | 20.6508 | 14.9854 | 10.424 | 6.419 | 4.5844 | 2.0478 | 1.7231 | 1.0895 | 0.24853 | -0.72612 | -1.5155 | -1.7416 | -0.90402 | -1.7686 | -3.3264 | -5.8293 | -7.2866 | -7.8542 | -9.156 | -8.6691 | -7.9245 | -5.7088 | -1.9398 | 5.0545 | 10.4138 | 19.343 | 30.6895 | 41.8347 | 53.2138 | 63.2046 | 68.9839 | 72.5562 | 71.6766 | 67.196 | 59.7808 | 52.0169 | 43.8069 | 36.4335 | 793.9933
Error % | -10.0859 | -11.9395 | -14.9708 | -21.7681 | -53.8261 | 59.395 | 15.451 | 7.4823 | 4.256 | 2.4146 | 1.4251 | 0.78957 | 0.53035 | 0.23026 | 0.19269 | 0.12319 | 0.028677 | -0.085859 | -0.18383 | -0.21639 | -0.11458 | -0.228 | -0.4347 | -0.76855 | -0.96356 | -1.0358 | -1.1985 | -1.1232 | -1.0144 | -0.7218 | -0.24306 | 0.63024 | 1.3088 | 2.4837 | 4.1047 | 5.9804 | 8.3763 | 11.3616 | 14.7986 | 19.4923 | 25.6841 | 34.8635 | 51.0773 | 93.9441 | 578.691 | -133.7011 | 793.9933
Sum of Errors = 3.3%

## Slide 35: Observations
Model is sufficient to investigate stray field reductions options.
Improvements to match measurements to Model include:
  More Tosca simulations on coil geometry.
  Increasing coil end lengths, widths.
  Try to add Q1 yoke to model.
  Accurate probe survey.

## Slide 36: Stray Field Data Points Midplane and Off Midplane (+/- 50mm)

## Slide 37: Below Mid-plane Setup
Probe is moved to lower hole
Probe holder is rotated about Vertical axis.

## Slide 38: 3500A Probe Rotated @ Midplane
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -401.7677 | -457.2238 | -510.6065 | -557.0736 | -584.15 | -577.7229 | -520.199 | -392.144 | -185.6291 | 90.8433 | 407.1555 | 718.7563 | 980.6999 | 1164.1684 | 1265.2976 | 1297.2877 | 1280.5435 | 1237.3755 | 1182.5035 | 974.2021 | 956.226 | 945.632 | 942.4939 | 944.6391 | 955.5025 | 972.2642 | 995.4192 | 1022.5146 | 1049.5343 | 1075.294 | 1090.5271 | 1085.3657 | 1052.7202 | 982.6027 | 872.4715 | 725.1436 | 553.0677 | 373.0468 | 203.1483 | 57.8295 | -54.7267 | -133.2802 | -181.4322
Measured | -383.25 | -435.7 | -486.45 | -530.48 | -554.45 | -541.99 | -473.18 | -330.31 | -110.3 | 183.04 | 513.25 | 831.14 | 1091.89 | 1270.03 | 1360.64 | 1383.29 | 1355.6 | 1303.96 | 1240.25 | 1023.36 | 998.41 | 977.82 | 967.87 | 967.62 | 975.97 | 991.1 | 1014.32 | 1041.49 | 1073.52 | 1105.13 | 1133.95 | 1140.54 | 1098.95 | 1016.04 | 889.93 | 730.11 | 549.83 | 366.77 | 200.08 | 61.78 | 5.84 | -114.75
Difference | 18.5177 | 21.5238 | 24.1565 | 26.5936 | 29.7 | 35.7329 | 47.019 | 61.834 | 75.3291 | 92.1967 | 106.0945 | 112.3837 | 111.1901 | 105.8616 | 95.3424 | 86.0023 | 75.0565 | 66.5845 | 57.7465 | 49.1579 | 42.184 | 32.188 | 25.3761 | 22.9809 | 20.4675 | 18.8358 | 18.9008 | 18.9754 | 23.9857 | 29.836 | 43.4229 | 87.8198 | 116.3473 | 143.5685 | 164.7864 | 177.0423 | 176.7832 | 163.6217 | -0.71097 | -1.8858 | -23.822 | 0.58111
Error % | -4.8318 | -4.94 | -4.9659 | -5.0131 | -5.3567 | -6.5929 | -9.9368 | -18.72 | -68.2947 | 50.3697 | 20.6711 | 13.5216 | 10.1833 | 8.3354 | 7.0072 | 6.2172 | 5.5368 | 5.1063 | 4.656 | 4.8036 | 4.2251 | 3.2918 | 2.6219 | 2.375 | 2.0971 | 1.9005 | 1.8634 | 1.8219 | 2.2343 | 2.6998 | 3.8293 | 7.6998 | 10.5871 | 14.1302 | 18.5168 | 24.2487 | 32.1523 | 44.6115 | -142.2505 | -116.5067 | -139.1202 | -66.6822
Sum of Errors = 10.6%

## Slide 39: 3500A Probe rotated @ 50 mm below midplane
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -367.0613 | -416.033 | -462.7141 | -501.1833 | -520.5531 | -503.5314 | -433.7031 | -295.2499 | -78.9147 | 192.0127 | 483.5376 | 747.1906 | 955.9455 | 1083.7098 | 1145.9682 | 1148.8984 | 1119.1725 | 1073.2305 | 1019.0609 | 833.67 | 819.1369 | 810.1262 | 805.4247 | 809.4476 | 818.7092 | 833.5972 | 852.49 | 876.3722 | 905.1635 | 929.0737 | 944.3072 | 947.9793 | 927.2643 | 873.8653 | 785.445 | 662.7308 | 516.438 | 355.4811 | 200.9234 | 65.4073 | -40.0868 | -114.3887 | 0
Measured | -360.99 | -410.88 | -459.61 | -499.13 | -519.34 | -502 | -425.85 | -274.81 | -44.31 | 246.7 | 557.33 | 842.17 | 1061.88 | 1206.84 | 1270.61 | 1275.6 | 1238.33 | 1180.73 | 1116.19 | 853.45 | 852.51 | 855.75 | 864.69 | 880.14 | 901.63 | 929.34 | 961.32 | 993.31 | 1023.22 | 1042.9 | 1040.34 | 1008.32 | 940.95 | 829.96 | 686.4 | 521.27 | 349.51 | 190.54 | 56.61 | -45.39 | -114.45
Difference | 6.0713 | 5.153 | 3.1041 | 2.0533 | 1.2131 | 1.5314 | 7.8531 | 20.4399 | 34.6047 | 54.6873 | 73.7924 | 94.9794 | 105.9345 | 123.1302 | 124.6418 | 126.7016 | 119.1575 | 107.4995 | 97.1291 | 43.3238 | 47.0853 | 46.3024 | 45.9808 | 46.5428 | 49.14 | 52.9678 | 56.1565 | 64.2363 | 78.9128 | 94.9207 | 113.0757 | 134.4547 | 155.505 | 167.2292 | 169.962 | 165.7889 | 148.5866 | -0.71097 | -1.8858 | -23.822 | 0.58111
Error % | -1.6818 | -1.2541 | -0.67537 | -0.41137 | -0.23359 | -0.30505 | -1.8441 | -7.4378 | -78.0969 | 22.1675 | 13.2403 | 11.2779 | 9.9761 | 10.2027 | 9.8096 | 9.9327 | 9.6224 | 9.1045 | 8.7018 | 5.0763 | 5.5231 | 5.4107 | 5.3176 | 5.2881 | 5.4501 | 5.6995 | 5.8416 | 6.4669 | 7.7122 | 9.1016 | 10.8691 | 13.3345 | 16.5264 | 20.1491 | 24.7614 | 31.8048 | 42.5128 | -142.2505 | -116.5067 | -139.1202 | -66.6822

## Slide 40: 3500A Probe rotated @ 50 mm above midplane
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Tosca | -367.0613 | -416.033 | -462.7141 | -501.1833 | -520.5531 | -503.5314 | -433.7031 | -295.2499 | -78.9147 | 192.0127 | 483.5376 | 747.1906 | 955.9455 | 1083.7098 | 1145.9682 | 1148.8984 | 1119.1725 | 1073.2305 | 1019.0609 | 833.67 | 819.1369 | 810.1262 | 805.4247 | 809.4476 | 818.7092 | 833.5972 | 852.49 | 876.3722 | 905.1635 | 929.0737 | 944.3072 | 947.9793 | 927.2643 | 873.8653 | 785.445 | 662.7308 | 516.438 | 355.4811 | 200.9234 | 65.4073 | -40.0868 | -114.3887 | 0
Measured | -361.01 | -399.85 | -455.15 | -475.31 | -514.57 | -514.56 | -498.9 | -275.72 | -51.37 | 555.36 | 841.87 | 1065.67 | 1205.15 | 1267.91 | 1260.66 | 1222.11 | 1176.43 | 1114.02 | 861.47 | 857.19 | 858.4 | 864.53 | 877.66 | 900.53 | 928.27 | 959.24 | 995.08 | 1020.01 | 1037.11 | 1032.63 | 999.03 | 930.69 | 815.63 | 673.43 | 500.2 | 342.72 | 187.48 | 59.42 | -36.14 | -105.36
Difference | 6.0513 | 16.183 | 7.5641 | 25.8733 | 5.9831 | -11.0286 | -65.1969 | 19.5299 | 27.5447 | 71.8224 | 94.6794 | 109.7245 | 121.4402 | 121.9418 | 111.7616 | 102.9375 | 103.1995 | 94.9591 | 51.3438 | 51.7653 | 48.9524 | 45.8208 | 44.0628 | 48.04 | 51.8978 | 54.0765 | 66.0063 | 75.7028 | 89.1307 | 105.3657 | 125.1647 | 145.245 | 152.8992 | 156.992 | 144.7189 | 141.7966 | -0.71097 | -1.8858 | -23.822 | 0.58111
Error % | -1.6762 | -4.0473 | -1.6619 | -5.4435 | -1.1627 | 2.1433 | 13.0681 | -7.0832 | -53.6203 | 12.9326 | 11.2463 | 10.2963 | 10.0768 | 9.6175 | 8.8653 | 8.4229 | 8.7723 | 8.524 | 5.96 | 6.039 | 5.7027 | 5.3001 | 5.0205 | 5.3346 | 5.5908 | 5.6374 | 6.6333 | 7.4218 | 8.5941 | 10.2036 | 12.5286 | 15.6062 | 18.7462 | 23.3123 | 28.9322 | 41.3739 | -142.2505 | -116.5067 | -139.1202 | -66.6822

## Slide 41: 3500A Midplane Symmetry
Chart: Row #1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
Measured Minus | -360.99 | -410.88 | -459.61 | -499.13 | -519.34 | -502 | -425.85 | -274.81 | -44.31 | 246.7 | 557.33 | 842.17 | 1061.88 | 1206.84 | 1270.61 | 1275.6 | 1238.33 | 1180.73 | 1116.19 | 853.45 | 852.51 | 855.75 | 864.69 | 880.14 | 901.63 | 929.34 | 961.32 | 993.31 | 1023.22 | 1042.9 | 1040.34 | 1008.32 | 940.95 | 829.96 | 686.4 | 521.27 | 349.51 | 190.54 | 56.61 | -45.39 | -114.45
Measured Plus | -361.01 | -399.85 | -455.15 | -475.31 | -514.57 | -514.56 | -498.9 | -275.72 | -51.37 | 555.36 | 841.87 | 1065.67 | 1205.15 | 1267.91 | 1260.66 | 1222.11 | 1176.43 | 1114.02 | 861.47 | 857.19 | 858.4 | 864.53 | 877.66 | 900.53 | 928.27 | 959.24 | 995.08 | 1020.01 | 1037.11 | 1032.63 | 999.03 | 930.69 | 815.63 | 673.43 | 500.2 | 342.72 | 187.48 | 59.42 | -36.14 | -105.36
Diff +Y - -Y | -0.02 | 11.03 | 4.46 | 23.82 | 4.77 | -12.56 | -73.05 | -0.91 | -7.06 | -1.97 | -0.3 | 3.79 | -1.69 | -2.7 | -14.94 | -16.22 | -4.3 | -2.17 | 8.02 | 4.68 | 2.65 | -0.16 | -2.48 | -1.1 | -1.07 | -2.08 | 1.77 | -3.21 | -5.79 | -7.71 | -9.29 | -10.26 | -14.33 | -12.97 | -21.07 | -6.79 | -3.06 | 2.81 | 9.25 | 9.09
Good agreement between upper and lower measurements +/- 20 Gauss or less. No asymmetry noted.

## Slide 42: Integral Stray By Field along Beamline @ 5.5° from Tosca Model
Tosca Model predicts an Integral By along beam line of 0.126 T.m at 4,000A
Criteria is 0.020T.m to prevent off beam dump window excursion.
Row 1 of mapping table vs true 5.5° beamline
Cyan = Row 1
Red = Beam Line

## Slide 43: Tosca Model - Mods to Yoke
Additional Iron pieces were included in the Tosca model.
Shapes and sizes are based on what geometry will allow.
Top addition may be hardest to add.
Various yoke configurations, using a +3mm coil offset were modeled.

## Slide 44: Adding Yoke Pieces to HB
As Built
All Pieces
Beam Line and Back Side
Beam Line

## Slide 45: Stray Field Along Beam Axis
All = 0.0705 T.m
Beam Side Only = 0.0772 T.m
Beam + Back Side = 0.0761 T.m
As Built = 0.1260 T.m
(see Slide 3)

## Slide 46: Conclusion
Tosca Model vs Measurement agreement is better than 1% inside the bore over the whole excitation range. Stray field predictions agree to within a few percent along the yoke. Larger errors at the front and back of the magnet (coil ends).
Errors in modeling the stray field are possible due to matching the as built coil end geometry and not modeling Q1’s Yoke.
Moving the coils +3mm in X improved the agreement between Tosca and measured data along the yoke region at higher excitation. At lower currents, moving the coils +3mm in X lowered the agreement, suggesting that there is a coil motion during excitation.
No asymmetry was observed in the off midplane measurements.
Additional Yoke pieces on beam side face will reduce the stray field by 39%. A factor of 4 reduction in stray field is still needed.

## Slide 47: HB @ 4000A - Target