## Slide 1: Harmonic Analysis of SHMS HB
Eric Sun
Steve Lassiter
Mike Fowler
March 31, 2015

## Slide 2: Hall Probe and Error Plot
Chart: Magnetic Field (kG)
Sensor model: LakeShore XHGT-9060
Sensor Type: InAs Hall Generator
Serial number: 55937

## Slide 3: Hysteresis Curve
Chart: Current (A)
| 0.10000009999999999 | 100.1001 | 200.00020000000001 | 300.05029999999999 | 399.80040000000002 | 499.95049999999998 | 1000.001 | 1500.1010000000001 | 1999.8520000000001 | 2200.1019999999999 | 2399.9520000000002 | 2599.9029999999998 | 2799.8029999999999 | 2999.9029999999998 | 3199.703 | 3399.654 | 3599.6039999999998 | 3799.5039999999999 | 3899.3539999999998
B (Pos, Up) [kG] | 0.003186 | 0.67293 | 1.3435 | 2.0146 | 2.6859 | 3.3575 | 6.716 | 10.0642 | 13.4017 | 14.7313 | 16.0544 | 17.3666 | 18.6657 | 19.948 | 21.2083 | 22.4344 | 23.6122 | 24.7502 | 25.3085
B (Neg, Up) [kG] | -0.0010998 | -0.67067 | -1.341 | -2.0115 | -2.6822 | -3.3533 | -6.7075 | -10.0561 | -13.3948 | -14.7246 | -16.0474 | -17.3602 | -18.6594 | -19.9425 | -21.204 | -22.4319 | -23.6118 | -24.7523 | -25.312
B (Pos, Down) [kG] | 24.7538 | 23.6197 | 22.4447 | 21.2201 | 19.9603 | 18.6782 | 17.3793 | 16.0664 | 14.7424 | 13.4114 | 10.0711 | 6.7212 | 3.3623 | 2.6905 | 2.0189 | 1.3472 | 0.67552 | 0.004248
B (Neg, Down) [kG] | -24.7567 | -23.6214 | -22.4454 | -21.2187 | -19.9582 | -18.6751 | -17.3758 | -16.0632 | -14.7381 | -13.4067 | -10.0664 | -6.7161 | -3.3602 | -2.6888 | -2.0179 | -1.3471 | -0.67633 | -0.0061133
Relative difference (Neg) [%] | 0.84426 | 0.45662 | 0.31898 | 0.24794 | 0.20441 | 0.12956 | 0.10262 | 0.088772 | 0.09177 | 0.098497 | 0.089847 | 0.084313 | 0.078546 | 0.069428 | 0.060049 | 0.040686 | 0.018091 | 0
Relative difference (Pos) [%] | 0.38473 | 0.27011 | 0.21404 | 0.17308 | 0.14141 | 0.078635 | 0.068661 | 0.072325 | 0.075264 | 0.074836 | 0.073606 | 0.066631 | 0.061464 | 0.055409 | 0.045743 | 0.031925 | 0.014367 | 0
Max relative difference (at -100 A) = 0.84%
Relative difference (at -3800 A) = 0.018%

## Slide 4: Upward Difference of Hysteresis Curve
Chart: Current (A)
| 0.10000009999999999 | 100.1001 | 200.00020000000001 | 300.05029999999999 | 399.80040000000002 | 499.95049999999998 | 1000.001 | 1500.1010000000001 | 1999.8520000000001 | 2200.1019999999999 | 2399.9520000000002 | 2599.9029999999998 | 2799.8029999999999 | 2999.9029999999998 | 3199.703 | 3399.654 | 3599.6039999999998 | 3799.5039999999999 | 3899.3539999999998
Upward difference (Pos - Neg) [kG] | 0.0020862 | 0.0022681 | 0.002548 | 0.0030916 | 0.0036698 | 0.0042132 | 0.0085042 | 0.0080821 | 0.0068952 | 0.0067552 | 0.0070412 | 0.0063835 | 0.0063589 | 0.0055127 | 0.0043915 | 0.0024583 | 0.00037043 | -0.0020378 | -0.0035479
Upward difference (Pos -Neg) [%] | 0.33705 | 0.18965 | 0.15346 | 0.13664 | 0.12548 | 0.12663 | 0.080306 | 0.05145 | 0.045856 | 0.043858 | 0.036758 | 0.034067 | 0.027635 | 0.020706 | 0.010958 | 0.0015688 | 0.0082333 | 0.014018
Max upward difference (at 100 A) = 0.34%
Upward difference (3900 A) = 0.014%

## Slide 5: Mapping Fixture
0°
Twenty points were measured on a 5.908 cm circle.
90°

## Slide 6: 1200 A (negative) at Location 27 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured [27] (error corrected) | -7.9607 | -7.9623 | -8.0052 | -8.0457 | -8.0741 | -8.0759 | -8.0701 | -8.0394 | -7.9785 | -7.9484 | -7.9387 | -7.9869 | -8.0782 | -8.1494 | -8.178 | -8.1645 | -8.1981 | -8.1775 | -8.0988 | -8.0169 | -7.9607
Tosca [27] | -7.9363 | -7.9445 | -7.9917 | -8.0609 | -8.0965 | -8.1048 | -8.0954 | -8.0598 | -8.0055 | -7.95 | -7.9466 | -7.9889 | -8.0799 | -8.1753 | -8.2098 | -8.2147 | -8.212 | -8.1709 | -8.0728 | -7.9796 | -7.9363
Difference (%) | -0.30746 | -0.22376 | -0.16775 | 0.18819 | 0.27717 | 0.35654 | 0.31346 | 0.25241 | 0.33742 | 0.019762 | 0.099189 | 0.02578 | 0.020599 | 0.31696 | 0.38669 | 0.6112 | 0.17022 | -0.080746 | -0.32168 | -0.46717 | -0.30746
max difference = 0.61% average difference = 0.09% median difference = 0.13%

## Slide 7: 2000 A (negative) at Location 27 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured [27] (error corrected) | -13.2475 | -13.2508 | -13.3197 | -13.3861 | -13.4362 | -13.439 | -13.4274 | -13.3761 | -13.2743 | -13.2252 | -13.2085 | -13.2849 | -13.4425 | -13.5585 | -13.6138 | -13.5911 | -13.651 | -13.6095 | -13.4791 | -13.3396 | -13.2475
Tosca [27] | -13.2205 | -13.2339 | -13.3121 | -13.4265 | -13.4855 | -13.4992 | -13.4837 | -13.4247 | -13.3349 | -13.2429 | -13.2376 | -13.3082 | -13.4598 | -13.6187 | -13.6762 | -13.6844 | -13.68 | -13.6114 | -13.448 | -13.2927 | -13.2205
max difference = 0.68% average difference = 0.19% median difference = 0.22%
Abnormal value is due to welds on inner tube of helium vessel.

## Slide 8: 3000 A (negative) at Location (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (27) [error corrected] | -19.728 | -19.7216 | -19.8189 | -19.9179 | -19.9815 | -19.9875 | -19.9739 | -19.9006 | -19.7493 | -19.6836 | -19.6614 | -19.7914 | -20.036 | -20.2176 | -20.302 | -20.2795 | -20.3593 | -20.2985 | -20.0922 | -19.8786 | -19.728
Tosca (27) | -19.7078 | -19.7205 | -19.8318 | -19.9972 | -20.0823 | -20.1014 | -20.0796 | -19.9944 | -19.8646 | -19.7339 | -19.7328 | -19.8459 | -20.0803 | -20.3238 | -20.4127 | -20.427 | -20.4188 | -20.3133 | -20.0628 | -19.8227 | -19.7078
max difference = 0.72% average difference = 0.29% median difference = 0.33%

## Slide 9: 3900 A (negative) at Location 27 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (27) [corrected] | -25.0077 | -24.9693 | -25.0872 | -25.2095 | -25.3036 | -25.313 | -25.2972 | -25.1979 | -25.0067 | -24.9244 | -24.9206 | -25.1309 | -25.5011 | -25.7855 | -25.9302 | -25.9162 | -26.007 | -25.8857 | -25.5752 | -25.2482 | -25.0077
Tosca (27) | -24.9988 | -24.9929 | -25.1334 | -25.3494 | -25.4642 | -25.4914 | -25.4607 | -25.3463 | -25.1765 | -25.0129 | -25.0341 | -25.2163 | -25.5671 | -25.9258 | -26.0651 | -26.0963 | -26.0737 | -25.9107 | -25.5423 | -25.1829 | -24.9988
max difference = 0.70% average difference = 0.36% median difference = 0.40%

## Slide 10: 3900 A (Positive) at Location 27 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured By (27) [corrected] | 25.0141 | 24.9796 | 25.0937 | 25.2155 | 25.315 | 25.3157 | 25.3043 | 25.206 | 25.0063 | 24.9315 | 24.9269 | 25.1269 | 25.5015 | 25.7942 | 25.9311 | 25.9143 | 26.0118 | 25.8895 | 25.5815 | 25.2548 | 25.0141
Tosca (27) | 24.9988 | 24.9929 | 25.1334 | 25.3494 | 25.4642 | 25.4914 | 25.4607 | 25.3463 | 25.1765 | 25.0129 | 25.0341 | 25.2163 | 25.5671 | 25.9258 | 26.0651 | 26.0963 | 26.0737 | 25.9107 | 25.5423 | 25.1829 | 24.9988
max difference = 0.70% average difference = 0.34% median difference = 0.39%

## Slide 11: 3900 A (Positive) at Location 26 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured B (26) [corrected] | 25.0042 | 24.9657 | 25.0836 | 25.2068 | 25.3052 | 25.3037 | 25.2912 | 25.1909 | 25.0017 | 24.9087 | 24.9136 | 25.1222 | 25.4965 | 25.766 | 25.9261 | 25.8902 | 25.9906 | 25.8853 | 25.5811 | 25.2523 | 25.0042
Tosca (26) | 24.9927 | 24.995 | 25.1362 | 25.3589 | 25.472 | 25.4953 | 25.4683 | 25.355 | 25.1795 | 25.0141 | 25.0276 | 25.2065 | 25.5547 | 25.9111 | 26.0507 | 26.0804 | 26.0594 | 25.8979 | 25.5305 | 25.1739 | 24.9927
max difference = 0.75% average difference = 0.37% median difference = 0.44%

## Slide 12: 3900 A (Positive) at Location 28 (measured on 3/11/2016)
Chart: Degree (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured B (28) [corrected] | 24.9901 | 24.9529 | 25.0688 | 25.1958 | 25.2889 | 25.2931 | 25.2829 | 25.1796 | 24.9828 | 24.9109 | 24.9074 | 25.125 | 25.4838 | 25.7601 | 25.92 | 25.9054 | 25.9926 | 25.8641 | 25.558 | 25.2327 | 24.9901
Tosca (28) | 24.9747 | 24.9616 | 25.0948 | 25.3074 | 25.4206 | 25.452 | 25.4174 | 25.3048 | 25.137 | 24.9813 | 25.0104 | 25.1982 | 25.5498 | 25.9105 | 26.0503 | 26.0822 | 26.0597 | 25.8957 | 25.5254 | 25.1639 | 24.9747
max difference = 0.68% average difference = 0.31% median difference = 0.35%
Abnormal value is due to welds on inner tube of helium vessel.

## Slide 13: Harmonics (3900 A, Positive, Location 27)
Measured (21 data points) |  |  |  |  | 
Order | Sine | Cosine | Amplitude | Phase | Difference (%)
n | A_n | B_n |  |  | 100*(A_i_T - A_i_M)/B_0_M
0 | 0.00 | -25365.71 | 25365.71 | 180.00 | 
1 | 360.18 | -47.99 | 363.37 | -97.59 | 0.15
2 | 16.26 | 341.01 | 341.40 | -2.73 | -0.50
3 | 32.89 | 3.46 | 33.07 | -84.00 | 0.00
4 | -4.54 | 60.68 | 60.85 | 4.28 | -0.22
5 | -19.29 | -1.12 | 19.32 | 93.31 | -0.11
6 | 9.66 | -15.83 | 18.55 | -148.61 | 0.06
7 | 1.54 | 4.22 | 4.50 | -20.07 | -0.01
8 | 9.44 | 12.25 | 15.47 | -37.62 | 0.05
9 | -7.15 | -2.19 | 7.48 | 107.02 | -0.03
Tosca (21 data points) |  |  |  |  | 
Order | Sine | Cosine | Amplitude | Phase | Difference (%)
n | A_n | B_n |  |  | 100*(B_i_T - B_i_M)/B_0_M
0 | 0.00 | -25471.04 | 25471.04 | 180.00 | 0.41
1 | 321.84 | -93.31 | 335.09 | -106.17 | 0.18
2 | 142.95 | 364.74 | 391.76 | -21.40 | -0.09
3 | 32.02 | 16.41 | 35.98 | -62.87 | -0.05
4 | 50.21 | 49.51 | 70.51 | -45.40 | 0.04
5 | 8.62 | 23.66 | 25.18 | -20.03 | -0.10
6 | -4.51 | 2.78 | 5.30 | 58.40 | -0.07
7 | 5.33 | 3.48 | 6.36 | -56.87 | 0.00
8 | -3.53 | 8.72 | 9.41 | 22.06 | 0.01
9 | 1.25 | 4.19 | 4.37 | -16.65 | -0.03
Tosca values were obtained using the same coordinates as the measured ones. The same Fourier Fit was employed for both cases.

## Slide 14: 3900 A (Positive) at Location 12 (measured on 3/23/2016)
Max/min difference = -5.8% average difference = -1.9% median difference = -1.7%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (12) [error corrected] | 11.4625 | 11.4148 | 11.5073 | 11.5716 | 11.4822 | 11.3405 | 11.4678 | 11.5169 | 11.4632 | 11.3696 | 11.3118 | 11.5588 | 11.8014 | 11.9479 | 11.5899 | 11.1884 | 11.3583 | 11.7294 | 11.7899 | 11.6698 | 11.4625
Tosca (12) | 11.3023 | 11.3494 | 11.4371 | 11.3421 | 10.9976 | 10.7203 | 10.9846 | 11.326 | 11.4396 | 11.3573 | 11.3235 | 11.3856 | 11.5082 | 11.5085 | 11.3268 | 11.1612 | 11.3305 | 11.5251 | 11.5055 | 11.3731 | 11.3023

## Slide 15: 3900 A (Positive) at Location 15 (measured on 3/23/2016)
Max/min difference = -3.7% average difference = -1.0% median difference = -0.88%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (15) [error corrected] | 18.0673 | 18.0625 | 18.2096 | 18.399 | 18.4557 | 18.4573 | 18.4402 | 18.2949 | 18.0056 | 17.8887 | 17.8908 | 18.2068 | 18.7571 | 19.2386 | 19.4139 | 19.309 | 19.4422 | 19.3336 | 18.9814 | 18.4899 | 18.0673
Tosca (15) | 17.9043 | 18.0216 | 18.2986 | 18.577 | 18.5611 | 18.4102 | 18.5382 | 18.5261 | 18.3432 | 18.0173 | 17.9156 | 18.0483 | 18.38 | 18.6879 | 18.7337 | 18.7046 | 18.75 | 18.7015 | 18.3747 | 18.0343 | 17.9043

## Slide 16: 3900 A (Positive) at Location 21 (measured on 3/23/2016)
Max/min difference = -1.0% average difference = -0.28% median difference = -2.6%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (21) [error corrected] | 24.314 | 24.2503 | 24.3324 | 24.4588 | 24.5084 | 24.5024 | 24.491 | 24.4257 | 24.2456 | 24.1676 | 24.1758 | 24.3904 | 24.8156 | 25.0816 | 25.1988 | 25.1646 | 25.29 | 25.2306 | 24.9389 | 24.5889 | 24.314
Tosca (21) | 24.2634 | 24.2778 | 24.4396 | 24.6584 | 24.748 | 24.7395 | 24.7405 | 24.6444 | 24.4779 | 24.2883 | 24.2867 | 24.4533 | 24.7852 | 25.1252 | 25.2576 | 25.2876 | 25.2683 | 25.1193 | 24.7698 | 24.4291 | 24.2634

## Slide 17: 3900 A (Positive) at Location 27 (measured on 3/23/2016)
Max/min difference = 0.71% average difference = 0.34% median difference = 0.36%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (27) [error corrected] | 25.0101 | 24.9724 | 25.0852 | 25.2017 | 25.3019 | 25.3093 | 25.2978 | 25.1995 | 24.9996 | 24.9249 | 24.9176 | 25.1266 | 25.5004 | 25.7731 | 25.9254 | 25.9397 | 26.006 | 25.8839 | 25.5767 | 25.2485 | 25.0101
Tosca (27) | 24.9988 | 24.9929 | 25.1334 | 25.3494 | 25.4642 | 25.4914 | 25.4607 | 25.3463 | 25.1765 | 25.0129 | 25.0341 | 25.2163 | 25.5671 | 25.9258 | 26.0651 | 26.0963 | 26.0737 | 25.9107 | 25.5423 | 25.1829 | 24.9988

## Slide 18: 3900 A (Positive) at Location 33 (measured on 3/23/2016)
Max/min difference = 1.1% average difference = 0.58% median difference = 0.62%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (33) [error corrected] | 24.2087 | 24.1405 | 24.2049 | 24.3066 | 24.3719 | 24.376 | 24.3706 | 24.2997 | 24.1546 | 24.1119 | 24.1727 | 24.3769 | 24.7352 | 24.9954 | 25.1759 | 25.2096 | 25.2369 | 25.0812 | 24.7742 | 24.4549 | 24.2087
Tosca (33) | 24.2779 | 24.21 | 24.281 | 24.4338 | 24.5278 | 24.5582 | 24.5233 | 24.4311 | 24.3086 | 24.2258 | 24.3068 | 24.5321 | 24.9031 | 25.2694 | 25.409 | 25.4473 | 25.4227 | 25.2587 | 24.8786 | 24.4985 | 24.2779

## Slide 19: 3900 A (Positive) at Location 39 (measured on 3/23/2016)
Max/min difference = 1.6% average difference = 0.39% median difference = 0.53%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (39) [error corrected] | 19.2551 | 19.1708 | 19.2158 | 19.4016 | 19.4726 | 19.5122 | 19.4993 | 19.3913 | 19.2523 | 19.2263 | 19.2989 | 19.487 | 19.9102 | 20.2123 | 20.4889 | 20.5146 | 20.5139 | 20.2774 | 19.8883 | 19.5329 | 19.2551
Tosca (39) | 19.3586 | 19.2169 | 19.2243 | 19.3062 | 19.3841 | 19.4056 | 19.3748 | 19.2997 | 19.232 | 19.2241 | 19.3741 | 19.6777 | 20.1269 | 20.5313 | 20.6701 | 20.7172 | 20.6986 | 20.5336 | 20.1043 | 19.6473 | 19.3586

## Slide 20: 3900 A (Positive) at Location 42 (measured on 3/23/2016)
Max/min difference = -3.1% average difference = -1.3% median difference = -1.3%
Chart: Angle (°)
| 0 | 18 | 36 | 54 | 72 | 90 | 108 | 126 | 144 | 162 | 180 | 198 | 216 | 234 | 252 | 270 | 288 | 306 | 324 | 342 | 360
Measured (42) [error corrected] | 13.5844 | 13.6659 | 13.8256 | 13.8922 | 14.0831 | 14.0908 | 14.1356 | 14.0381 | 13.7816 | 13.7293 | 13.6854 | 13.853 | 14.1651 | 14.3713 | 14.4587 | 14.3767 | 14.4193 | 14.3658 | 14.0681 | 13.8043 | 13.5844
Tosca (42) | 13.567 | 13.4931 | 13.5465 | 13.6617 | 13.7151 | 13.7258 | 13.7126 | 13.6613 | 13.5712 | 13.5117 | 13.5992 | 13.8175 | 14.1214 | 14.3008 | 14.1994 | 14.087 | 14.2273 | 14.3236 | 14.1043 | 13.7849 | 13.567

## Slide 21: Summary
Max upward difference (Positive – Negative) of hysteresis curve is less than 0.34% (at 100 A); upward difference is only 0.018% at 3900 A.
Max difference of By field on a 5.908 cm circle is 0.72% at Location 27 (z = 15.5 mm), and 5.8% at Location 12 (z = 396.5 mm).
A_n and B_n of Fourier Fit of 20 measured points and 20 Tosca-calculated points are very close with all measured currents at Location 26, 27, and 28. B_n difference is less than 0.41% for all currents at Location 27.